## Instructions
Stage 4 adjudication — instructions
Worksheet: adjudication_worksheet.csv — 71 cases, 35 columns.
The worksheet holds two queues, deliberately mixed together:
21 DISAGREEMENT cases — every case where the production trio failed to reach a 2-of-3 verdict (15 all_three_disagree, 6 two_disagree_one_uncertain). This is the required Stage 4 queue.
50 CONVERGENT spot-check cases — a stratified sample of cases the trio agreed on, to measure how often you would overturn the ensemble where it was confident. The spec targets an override rate at or below 0.10.
You cannot tell which is which, and that is intentional. The rows are shuffled (seed 20260720) and carry no queue label. If you could see that a case was one the judges agreed on, you would unconsciously treat it as already settled, which is precisely the bias the spot-check exists to detect. The mapping lives in adjudication_sample_key.csv. Do not open that file until you have finished.
For the same reason the worksheet contains case evidence only: no judge verdicts, no judge reasoning. Your verdict is therefore independent by construction, which is what makes both the author-versus-judge agreement statistic and the override rate interpretable. Judge verdicts for the 21 disagreements are in adjudication_queue.csv, again for afterwards only.
Because the order is random, any prefix of the sheet is a valid random subset of both queues. If you want to split the work, doing rows 1 to 35 today and the rest later costs you nothing statistically.
How the sample was drawn
| Sampled | Population | Weight
CORRECT-DETECTION | 15 | 2677 | 0.5903
EXISTING-RULE-MISBEHAVIOR | 12 | 1161 | 0.256
GENERATOR-ARTIFACT | 8 | 398 | 0.0878
REAL-GAP | 12 | 289 | 0.0637
OUT-OF-SCOPE | 3 | 10 | 0.0022
The question you are answering
For each case: did the catalog behave correctly on this package? That decomposes into two sub-questions, in order.
Does the package actually instantiate the defeater it was asked to? Compare target_weakener against the package evidence in columns 13 to 29. If it does not, the verdict is GENERATOR-ARTIFACT regardless of what the engine did, because the catalog cannot be faulted for missing something that is not there.
If it does, did the engine respond correctly? Compare expected_rule and target_rule_fired against rules_fired.
Columns
What the generator was asked to build (3 to 8). target_weakener is the defeater it was told to instantiate, defeater_type is the D1/D2/D3 class, source_taxonomy is the literature origin, subtlety is how hidden it was told to make it. For the 7 gap_probe cases these read (none - gap_probe): those probe literature defeaters the catalog has no rule for, so nothing firing is the expected outcome and REAL-GAP is the live hypothesis.
What the engine did (9 to 12). engine_outcome is COV-HIT (target rule fired), COV-WRONG (target did not fire but something else did), COV-CLEAN-WRONG (a clean control that fired anyway), or GEN-INVALID (failed SHACL). rules_fired is everything that fired. In this queue: 16 COV-WRONG, 5 COV-HIT.
What the package says (13 to 29). This is your evidence.
package_description, cou_intended_use, cou_decision_consequence, cou_model_influence — what is being claimed and how much the model matters.
profile — ProfileMinimal or ProfileComplete. Minimal packages legitimately omit fields Complete requires, which matters when judging whether an omission is a defeater or just the profile.
decision, assurance_level, model_risk_level — the assessment outcome and how much rigor was required.
factors_detail — every credibility factor as factorType [status] req/ach AC=yes¦MISSING :: acceptance criteria. This is usually where the defeater lives.
Three pre-computed flags, since these are the common defeater shapes: factors_missing_acceptance_criteria (the W-AR-01 undermining shape), factors_not_assessed (the W-CON-01 shape), factors_below_required_level (achieved under required, an accepted-with-shortfall shape).
validation_results, declared_weakeners, decision_rationale.
package_in_bundle (30) — read the full package when the summary is not enough:
tar -xzf dev/build/adversarial/phase2/2026-04-26/judge_ready_bundle.tgz \
-O judge_ready_bundle/packages/<case_id>.jsonld ¦ jq .
What you fill in (31 to 35)
author_verdict — one of, spelled exactly: CORRECT-DETECTION, REAL-GAP, GENERATOR-ARTIFACT, EXISTING-RULE-MISBEHAVIOR, OUT-OF-SCOPE, UNCERTAIN.
Verdict | Use when
CORRECT-DETECTION | The package instantiates the defeater and the target rule caught it, for the right reason.
REAL-GAP | The package instantiates a genuine defeater the catalog has no rule for. This is the class that drives catalog extension, so hold it to the evidence.
GENERATOR-ARTIFACT | The package does not actually instantiate the intended defeater. The generator failed.
EXISTING-RULE-MISBEHAVIOR | A rule fired but wrongly: wrong rule, false positive, or right rule for the wrong reason.
OUT-OF-SCOPE | The bundle is structurally complete but lacks the evidence needed to evaluate the defeater. Say which evidence is missing.
UNCERTAIN | After inspecting the package, two or more verdicts each stay defensible and none dominates. A genuine tie, not a hard case you would rather not call.
author_section_6_7 — only when the verdict is REAL-GAP. Use a Tier-1 id if one fits (W-EV-01, W-EV-02, W-REQ-01, W-CX-01, W-AR-06, W-AR-07), otherwise name the pattern.
author_confidence — 0.0 to 1.0. Reflect real uncertainty rather than defaulting high.
author_rationale — a sentence or two on what decided it. A reviewer will read this, so write it for them.
author_notes — anything else.
One caveat worth carrying
These packages were generated on 2026-04-26 against the catalog before the Phase 2.5 refinements took the negative-control clean rate from 0% to 97.1%. A gap the earlier catalog missed is not necessarily a gap in the current one. If you mark something REAL-GAP and suspect v0.5.15.1 already closed it, say so in author_notes. This is the version-skew issue in PHASE3_STATUS_REPORT.md, and it is still open.
After you finish
Save as CSV in place. Corpus-wide agreement statistics:
uofa adversarial adjudicate \
--judgments-a dev/build/adversarial/phase3/production/run-1/judgments_A.jsonl \
--judgments-b dev/build/adversarial/phase3/production/run-1/judgments_B.jsonl \
--judgments-c dev/build/adversarial/phase3/production/run-1/judgments_C.jsonl \
--out dev/build/adversarial/phase3/adjudication/
That yields the three pairwise Cohen's kappa values, Fleiss' kappa, and per-pair confusion matrices.
Then open adjudication_sample_key.csv and join on case_id to get two readouts.
Author versus judge agreement, over the 21 disagreement cases: how your verdict relates to each judge's, which is what the disagreement queue was for.
Spot-check override rate, over the 50 convergent cases: the share where your verdict differs from ensemble_majority_verdict. Compute it per stratum, since the sample is stratified rather than proportional. For a population estimate, reweight by the stratum_weight column:
override_rate = Σ (stratum_weight × per-stratum override rate)
The spec target is 0.10 or lower. Note the precision honestly: 12 REAL-GAP cases can distinguish "rare" from "common" but cannot pin a rate to two decimal places, and with 3 of 10 OUT-OF-SCOPE cases that stratum is indicative only.
The REAL-GAP stratum is the one to read closely. Those 12 cases are the direct check on whether the 289 REAL-GAP verdicts underpinning the 6-of-6 Tier-1 result hold up under author review. If you overturn a meaningful share of them, that is a finding about the Tier-1 claim, not just about the ensemble.

## adjudication_worksheet
# | case_id | coverage_intent | target_weakener | defeater_type | source_taxonomy | subtlety | base_cou | engine_outcome | expected_rule | target_rule_fired | rules_fired | package_name | package_description | profile | decision | assurance_level | model_risk_level | cou_decision_consequence | cou_model_influence | cou_intended_use | n_factors | factors_missing_acceptance_criteria | factors_not_assessed | factors_below_required_level | factors_detail | validation_results | declared_weakeners | decision_rationale | package_in_bundle | author_verdict | author_section_6_7 | author_confidence | author_rationale | author_notes | guide_target_rule | guide_rules_fired | scan_bucket | scan_basis | scan_6_7_candidate
1 | adv-2026-p2-002-w-ar-02_medium_morrison-cou1-v09 | confirm_existing | W-AR-02 | D2 | jarzebowicz-wardzinski/argument_defeaters/D2-rebutting | medium | morrison-cou1 | COV-HIT | W-AR-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Discretization error(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial CFD solver with documented QA procedures and version control meets MRL 2 requirements. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes standard benchmark verification cases relevant to rotating machinery and viscous flow. - Discretization error [assessed] req3/ach2 AC=yes :: Grid convergence index (GCI) computed from at least three systematically refined meshes; GCI on primary QoI (hemolysis index) must… - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions used for validation are documented and traceable to the simulation boundary conditions. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted hemolysis index values compared against experimental measurements with quantified agreement metric. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs computed by the model (hemolysis index, wall shear stress distribution) are directly relevant to the clinical endpoint of pum… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities use a geometry and operating regime sufficiently similar to the COU device and conditions. - Numerical solver error [assessed] req1/ach1 AC=yes :: Iterative convergence demonstrated by residual reduction of at least 4 orders of magnitude for all transport equations. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; operator error controls addressed through standard SOP review. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; model form uncertainty acknowledged qualitatively. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; input parameters sourced from published literature with documented provenance. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; experimental sample variability addressed through use of multi-laboratory FDA benchmark dataset. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; parameter equivalency between validation case and COU addressed qualitatively. | Mesh convergence study for FDA generic centrifugal pump geometry; Relative hemolysis index comparison against experimental PIV and hemolysis measurements; SST k-omega turbulence model code verification against benchmark cases | (none declared) | The credibility assessment team reviewed all 13 ASME V&V 40 factors for Model Risk Level 2. Seven factors were formally assessed and six were scoped out as not required at MRL 2. The CFD model demonstrates satisfactory performance across the validation operating envelope, with mean absolute error of 12.3% on NIH and correct rank-ordering of hemolysis severity across all five operating points. The … | judge_ready_bundle/packages/adv-2026-p2-002-w-ar-02_medium_morrison-cou1-v09.jsonld | CORRECT-DETECTION |  | 1 | Offset rationale was in prose not as a structured property. |  | W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. [Argumentation] Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
2 | adv-2026-p2-113-faults-software_low-v02 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/faults-software | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Version-controlled solver with documented unit test suite passing all nominal test cases. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver verified against at least one analytical benchmark relevant to rotating internal flow. - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for the primary quantity of interest (NIH) on the production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions documented and matched to CFD boundary conditions within stated tolerances. - Output comparison [assessed] req2/ach2 AC=yes :: Mean absolute relative error in NIH < 15% across all 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (NIH and RH index) directly linked to clinical hemolysis endpoint. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments cover the COU operating envelope (2–6 L/min, 2000–3500 RPM) with the same pump geometry. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor; basic documentation of solver settings sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; analyst qualifications documented. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model and hemolysis damage model selection justified with supporting literature. - Model inputs [assessed] req2/ach2 AC=yes :: All material properties and boundary conditions traceable to documented sources with stated uncertainties. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; sample documentation sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; qualitative equivalency argument sufficient. | Grid convergence study — three-level mesh refinement; Bench-scale hemolysis index comparison — 5 operating points; Software quality assurance — unit and integration test suite; Model input sensitivity — blood viscosity and power-law parameters | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors were scoped out as not required at MRL 2 per the ASME V&V 40 framework. The CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions for the Class II CPB device. The relative hemolysis index RH < 1 was confirmed at all 5 operating points, satisfying t… | judge_ready_bundle/packages/adv-2026-p2-113-faults-software_low-v02.jsonld | OUT-OF-SCOPE |  | 1 | the weakness is expressed in prose only and also it is not clear what candidate rule is expected to fire here |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. | NONE - no Tier-1 candidate names this defeater
3 | adv-2026-p2-021-compound-01_medium_morrison-cou1-v04 | confirm_existing | COMPOUND-01 | structural | khakzad/argumentation/undercutting-d3 | medium | morrison-cou1 | COV-HIT | COMPOUND-01 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented verification tests and version-controlled source. Achieved level meets required level for MRL 2. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must reproduce manufactured solutions for Navier-Stokes equations to second-order accuracy. Achieved level meets required l… - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% at the finest mesh level used in production runs. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions for validation must match COU operating envelope within documented tolerances. Flow rate ±5%, rotatio… - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (MAE or RMSE) between CFD predictions and experimental measurements must be within 15% for the prim… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (relative hemolysis index RH) must be directly computable from CFD outputs and directly comparable to the clinical acc… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must use geometry, fluid properties, and operating conditions representative of the COU device and patient p… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. Iterative convergence monitoring sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. Peer review of simulation setup sufficient. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. Turbulence model selection rationale documented. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. Nominal input values with documented sources sufficient. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. Sample characterization documentation sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. Qualitative mapping between simulation inputs and experimental conditions sufficient. | Hemolysis bench-scale validation against FDA in-vitro data; Shear stress distribution validation against PIV measurements | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped to not-assessed per ASME V&V 40 Table 3 requirements for Class II devices at MRL 2. CFD model demonstrates sufficient credibility for identifying worst-case hemolysis operating conditions. Relative hemolysis index predictions within 15% acceptance band across all validation operating p… | judge_ready_bundle/packages/adv-2026-p2-021-compound-01_medium_morrison-cou1-v04.jsonld | CORRECT-DETECTION |  | 1 | at least one Critical and one High Level-1 fire |  | COMPOUND-01 (Critical) — Critical and High weakeners coexist. [Compound] Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. Wrong only if the underlying set doesn't support it. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
4 | adv-2026-p2-013-w-si-01_high_morrison-cou2-v13 | confirm_existing | W-SI-01 | structural | khakzad/argumentation/undermining-d1 | high | morrison-cou2 | GEN-INVALID | W-SI-01 | FALSE |  | Morrison COU2 — VAD hemolysis credibility assessment (W-SI-01 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. This package i… | ProfileMinimal | Not accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach2); Numerical code verification(req4>ach2); Discretization error(req4>ach3); Test conditions(req4>ach2); Output comparison(req4>ach2); Relevance of the quantities of interest(req4>ach3); Relevance of the validation activities to the COU(req4>ach2); Numerical solver error(req4>ach1); Use error(req3>ach1); Model form(req5>ach2); Model inputs(req4>ach1); Test samples(req3>ach1); Equivalency of input parameters(req4>ach1) | - Software quality assurance [assessed] req4/ach2 AC=yes :: CFD solver must have documented verification test suite, version control, and regression testing consistent with MRL 5 requirement… - Numerical code verification [assessed] req4/ach2 AC=yes :: Order-of-accuracy tests and method-of-manufactured-solutions (MMS) verification must be demonstrated for all governing equations u… - Discretization error [assessed] req4/ach3 AC=yes :: Formal grid convergence study using Richardson extrapolation must be performed across at least three systematically refined meshes… - Test conditions [assessed] req4/ach2 AC=yes :: Experimental test conditions used for validation must closely replicate VAD operating envelope including fluid temperature, hemato… - Output comparison [assessed] req4/ach2 AC=yes :: Quantitative comparison of predicted versus measured hemolysis index must include uncertainty bounds on both simulation and experi… - Relevance of the quantities of interest [assessed] req4/ach3 AC=yes :: The primary quantity of interest (absolute hemolysis index) must be directly measurable in validation experiments and directly com… - Relevance of the validation activities to the COU [assessed] req4/ach2 AC=yes :: Validation activities must cover the full VAD operating envelope (5 flow-speed combinations) with geometry identical or analytical… - Numerical solver error [not-assessed] req4/ach1 AC=yes :: Iterative convergence, round-off error, and solver-specific numerical diffusion must be quantified and shown to be negligible rela… - Use error [not-assessed] req3/ach1 AC=yes :: Evidence of independent review of model setup, boundary condition specification, and post-processing procedures must be provided t… - Model form [not-assessed] req5/ach2 AC=yes :: The power-law hemolysis model must be shown to be the appropriate model form for the flow regimes encountered in the VAD, includin… - Model inputs [not-assessed] req4/ach1 AC=yes :: Sensitivity analysis must demonstrate that uncertainty in blood rheological parameters (viscosity, yield stress, hematocrit) does … - Test samples [not-assessed] req3/ach1 AC=yes :: Experimental samples must be characterized for lot-to-lot variability in hemolytic susceptibility, and a minimum of three independ… - Equivalency of input parameters [not-assessed] req4/ach1 AC=yes :: Correspondence between simulation input parameters and experimental conditions must be formally established, including pump geomet… | Bench-scale hemolysis validation against FDA generic pump experimental data; Monte Carlo uncertainty quantification on absolute hemolysis index predictions | (none declared) | The credibility assessment package is accepted as a complete record of the MRL 5 evaluation for the VAD hemolysis COU. All 13 credibility factors have been evaluated or formally flagged as epistemic gaps. The package is submitted for regulatory review with the understanding that six factors remain unassessed and that all assessed factors fall below required levels, resulting in an overall Not Acce… | judge_ready_bundle/packages/adv-2026-p2-013-w-si-01_high_morrison-cou2-v13.jsonld | GENERATOR-ARTIFACT |  | 1 | signature field is missing but the SHACL probably failed the package not allowing the rule to fire |  | W-SI-01 (Medium) — No digital signature. [Structural] Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT. | (no rules fired)
5 | adv-2026-p2-011-w-on-01_low_nagaraja-cou1-v06 | confirm_existing | W-ON-01 | structural | khakzad/ontological/scope-drift | low | nagaraja-cou1 | GEN-INVALID | W-ON-01 | FALSE |  | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | [6 empty] | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach3 AC=yes :: Achieved level must meet or exceed required level for Medium model risk. - Numerical code verification [assessed] req2/ach2 AC=yes :: Achieved level must meet or exceed required level. - Discretization error [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Numerical solver error [assessed] req2/ach2 AC=yes :: Achieved level must meet or exceed required level. - Use error [assessed] req2/ach2 AC=yes :: Achieved level must meet or exceed required level. - Model form [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Model inputs [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Test samples [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Test conditions [assessed] req3/ach2 AC=yes :: Achieved level must meet or exceed required level. - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Output comparison [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Achieved level must meet or exceed required level. | F1717 compression-bending stiffness and yield load comparison — FEA vs. N=11 experimental constructs | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed their required levels. The single shortfall in Test conditions (achieved Gradation 2 vs. required Gradation 3) was reviewed by the team and deemed acceptable given the prescriptive single-condition structure of ASTM F1717 and the high performance on all other factors. Overall credibility is sufficient to support a Medium model risk class… | judge_ready_bundle/packages/adv-2026-p2-011-w-on-01_low_nagaraja-cou1-v06.jsonld | GENERATOR-ARTIFACT |  | 1 | hasContextOfUse missing but SHACL validation probably caught it preventing rule from firing |  | W-ON-01 (Critical) — No Context of Use at all. [Ontological] Look at: cou_* columns populated? If the package has any COU, firing is wrong. Note: flaw cannot exist in a schema-valid package (COU is sh:minCount mandatory), so a "clean-looking" case targeting this is likely GENERATOR-ARTIFACT (package failed validation upstream). | (no rules fired)
6 | adv-2026-p2-017-w-con-03_high_morrison-cou2-v20 | confirm_existing | W-CON-03 | structural | gohar/requirements/stale | high | morrison-cou2 | COV-WRONG | W-CON-03 | FALSE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (W-CON-03 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. This synthetic… | ProfileMinimal |  |  |  | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Discretization error(req4>ach3); Test conditions(req4>ach3); Output comparison(req4>ach3); Relevance of the quantities of interest(req4>ach3); Relevance of the validation activities to the COU(req4>ach3); Numerical solver error(req4>ach2); Use error(req4>ach2); Model form(req5>ach2); Model inputs(req4>ach2); Test samples(req4>ach2); Equivalency of input parameters(req4>ach2) | - Software quality assurance [assessed] req4/ach3 AC=yes :: Version-controlled solver with documented test suite covering all active physics modules; independent code review completed. - Numerical code verification [assessed] req4/ach3 AC=yes :: Order-of-accuracy verification on canonical rotating-flow benchmark with demonstrated second-order spatial convergence. - Discretization error [assessed] req4/ach3 AC=yes :: Grid convergence index (GCI) reported for all QoI at all 5 COU operating points; GCI < 5% on fine grid. - Test conditions [assessed] req4/ach3 AC=yes :: Experimental boundary conditions match COU operating envelope within ±5% for all flow-speed combinations; fluid temperature and vi… - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative agreement metric (e.g. E_rms) defined and computed for each QoI; statistical significance of agreement established. - Relevance of the quantities of interest [assessed] req4/ach3 AC=yes :: QoI directly maps to clinical hemolysis threshold metric; traceability from CFD output to regulatory acceptance criterion document… - Relevance of the validation activities to the COU [assessed] req4/ach3 AC=yes :: Validation activities performed on geometry, fluid, and operating conditions sufficiently similar to COU; similarity rationale doc… - Numerical solver error [not-assessed] req4/ach2 AC=yes :: Iterative convergence criteria defined and demonstrated; floating-point round-off assessed for all active solver modules. - Use error [not-assessed] req4/ach2 AC=yes :: Documented user qualification, simulation setup checklist, and independent review of model inputs prior to each simulation run. - Model form [not-assessed] req5/ach2 AC=yes :: Sensitivity analysis demonstrates that turbulence model selection and hemolysis constitutive law form do not dominate prediction u… - Model inputs [not-assessed] req4/ach2 AC=yes :: All model input parameters traced to measurement data or literature with documented uncertainty; Monte Carlo input distributions v… - Test samples [not-assessed] req4/ach2 AC=yes :: Sufficient replicate experiments (n ≥ 5 per operating point) with statistical power analysis; sample provenance and storage condit… - Equivalency of input parameters [not-assessed] req4/ach2 AC=yes :: Quantitative mapping between experimental and simulation input parameters demonstrated; any parameter substitutions (e.g. blood an… | PIV velocity field comparison — FDA generic centrifugal pump; Absolute hemolysis index bench validation — VAD operating envelope; Monte Carlo uncertainty quantification — hemolysis predictions | (none declared) | Despite multiple credibility factor gaps at MRL 5, the assessment team accepted the package on the basis that Monte Carlo UQ bounds remain below the clinical hemolysis threshold at all operating points and that the identified gaps are mitigated by conservative safety margins in the acceptance criterion. Residual risks are acknowledged and a follow-on assessment is planned prior to IDE submission. | judge_ready_bundle/packages/adv-2026-p2-017-w-con-03_high_morrison-cou2-v20.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | Uofa timestamp 2026-03-14T11:42:00Z whereas comparison valdation result has later timestamp 2026-03-14T11:47:23Z |  | W-CON-03 (High) — Evidence timestamp later than signature timestamp. [Consistency] Look at: does the package structurally carry BOTH timestamps? Rarely generated (0/180). Absent → GENERATOR-ARTIFACT. Present with evidence > signature → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
7 | adv-2026-p2-003-w-ar-03_high_nagaraja-cou1-v01 | confirm_existing | W-AR-03 | D3 | jarzebowicz-wardzinski/argument_defeaters/D3-undercutting | high | nagaraja-cou1 | COV-HIT | W-AR-03 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-03; W-AR-05; W-EP-01; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req2>ach1) | - Software quality assurance [assessed] req2/ach3 AC=yes :: FEA software (Ansys Mechanical 2022 R2) under documented SQA plan with version control, regression testing, and audit trail. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems with known analytical solutions demonstrate solver correctness for structural contact and large-displacement be… - Discretization error [assessed] req2/ach2 AC=yes :: Richardson extrapolation or GCI analysis demonstrates mesh-induced error <2% on primary QoIs. - Numerical solver error [assessed] req1/ach2 AC=yes :: Iterative solver residuals below 1e-6; energy balance verified for nonlinear load steps. - Use error [assessed] req2/ach2 AC=yes :: Independent peer review of model setup, boundary conditions, and material assignments by analyst not involved in original model co… - Model form [assessed] req3/ach3 AC=yes :: Continuum elastoplastic constitutive model with isotropic hardening demonstrated appropriate for Ti-6Al-4V pedicle screw under qua… - Model inputs [assessed] req3/ach3 AC=yes :: Material properties derived from lot-specific coupon testing; geometric inputs from CMM measurements of production specimens; fric… - Test samples [assessed] req2/ach2 AC=yes :: N≥10 specimens from production lots with documented dimensional and material traceability. - Test conditions [assessed] req2/ach1 AC=yes :: Testing performed under at least two distinct load configurations or environmental conditions to characterize model sensitivity. - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: FEA model inputs demonstrably derived from or traceable to measured specimen properties; no free parameters adjusted post-hoc. - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison metric E/uval < 1.96 and percent difference <10% for all primary QoIs; statistical uncertainty in experime… - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: QoIs (stiffness, yield force, force at 40 mm displacement) directly map to COU decision quantities and regulatory submission requi… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario geometry, loading, and boundary conditions match COU scenario within documented similarity bounds. | Construct stiffness QoI comparison — FEA vs. N=11 bench specimens; Yield force QoI comparison — FEA vs. N=11 bench specimens | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed required levels. The single shortfall (Test conditions, achieved 1 vs. required 2) is offset by the strong quantitative agreement across all three QoIs (E/uval < 1.96, percent difference < 10%) and the high relevance of validation scenario to COU. The validation dataset, assembled through retrospective review of archived laboratory recor… | judge_ready_bundle/packages/adv-2026-p2-003-w-ar-03_high_nagaraja-cou1-v01.jsonld | CORRECT-DETECTION |  | 1 | There is a requirement stating compression-bending validation needs to be "controlled-experiment" whereas the Validation result for it is of activityType "observational-study" |  | W-AR-03 (High) — Requirement demands verification method X; activity used activityType Y. [Argumentation] Look at: factors_detail / validation_results for a required method versus the activity's recorded type. Both fields must exist structurally. If activityType was never populated (common in generated packages), lean GENERATOR-ARTIFACT. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-03 (High) — Requirement demands verification method X; activity used activityType Y. Look at: factors_detail / validation_results for a required method versus the activity's recorded type. Both fields must exist structurally. If activityType was never populated (common in generated packages), lean GENERATOR-ARTIFACT. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
8 | adv-2026-p2-205-nc-clean-scoped-out-factors_medium_nagaraja-cou1-v01 | negative_control | (not recorded) |  | control/none | medium | nagaraja-cou1 | COV-WRONG | (not recorded) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-02; W-CON-01; W-EP-01; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | Numerical solver error; Relevance of the quantities of interest; Relevance of the validation activities to the COU | Numerical solver error(scoped-out); Relevance of the quantities of interest(scoped-out); Relevance of the validation activities to the COU(scoped-out) | (none) | - Software quality assurance [assessed] req2/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/sqa-level2 - Numerical code verification [assessed] req2/ach2 AC=yes :: https://uofa.net/nagaraja/criteria/code-verification-level2 - Discretization error [assessed] req3/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/discretization-level3 - Numerical solver error [scoped-out] reqNone/achNone AC=MISSING :: - Use error [assessed] req2/ach2 AC=yes :: https://uofa.net/nagaraja/criteria/use-error-level2 - Model form [assessed] req3/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/model-form-level3 - Model inputs [assessed] req3/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/model-inputs-level3 - Test samples [assessed] req3/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/test-samples-level3 - Test conditions [assessed] req2/ach2 AC=yes :: https://uofa.net/nagaraja/criteria/test-conditions-level2 - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: https://uofa.net/nagaraja/criteria/equiv-inputs-level3 - Output comparison [assessed] req3/ach4 AC=yes :: https://uofa.net/nagaraja/criteria/output-comparison-level3 - Relevance of the quantities of interest [scoped-out] reqNone/achNone AC=MISSING :: - Relevance of the validation activities to the COU [scoped-out] reqNone/achNone AC=MISSING :: | Construct stiffness FEA vs. experimental comparison; Yield force FEA vs. experimental comparison; Force at 40 mm displacement FEA vs. experimental comparison (Lab 1 and Lab 2) | (none declared) | The FEA model ensemble for the non-cannulated 7.5 mm pedicle screw construct under ASTM F1717 static compression-bending achieves the required credibility level for a Medium model risk / Medium decision consequence COU. All ten assessed factors meet or exceed their required gradation levels. Three factors (Numerical solver error, Relevance of QoIs, Relevance of validation activities to COU) are sc… | judge_ready_bundle/packages/adv-2026-p2-205-nc-clean-scoped-out-factors_medium_nagaraja-cou1-v01.jsonld | GENERATOR-ARTIFACT |  | 1 | package COU has neither applicability constraint nor operating envelope and W-ON-02 fires rightly |  | NEGATIVE CONTROL — clean package. Clean packages. Any firing beyond zero criticals is the question: is the fired-on absence genuinely present (then the package isn't clean and it's GENERATOR-ARTIFACT territory) or is the rule misreading (EXISTING-RULE-MISBEHAVIOR)? | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-02 (Medium) — UQ reported but no sensitivity analysis documented. Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. W-CON-01 (High) — Decision Accepted but a factor has neither requiredLevel nor achievedLevel. Look at: decision = Accepted AND any factor with both levels missing (factors_detail). Note the v0.5.14 guard: factors with status 'not-assessed' are exempt; if the only levelless factors are explicitly not-assessed, firing is wrong. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
9 | adv-2026-p2-115-configuration_low-v04 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/configuration | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (configuration gap variant) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. T… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial CFD solver with documented version control, regression test suite, and release notes. Version used in validation must b… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes at least one benchmark verification test relevant to rotating turbomachinery flow with scalar transport. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index reported for hemolysis index QOI; GCI < 5% on fine mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental bench conditions (flow rate, rotational speed, fluid properties) are documented and match the five CFD operating poin… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted NIH and RH values compared quantitatively against bench measurements; prediction error reported with uncertainty bounds. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QOIs (normalized index of hemolysis, relative hemolysis index) are directly linked to the clinical endpoint and the regulatory acc… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, fluid, and operating conditions are sufficiently similar to the COU device and intended operating envelope. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor when software quality assurance and code verification are assessed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for Class II device with medium model influence. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model and hemolysis power-law model selection are justified with reference to literature and sensitivity analysis. - Model inputs [assessed] req2/ach2 AC=yes :: All material properties, boundary conditions, and geometric parameters are documented with sources and uncertainty estimates. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 when test conditions and output comparison factors are assessed. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this COU given that the validation geometry and operating conditions directly match the COU. | Mesh sensitivity study — hemolysis index convergence; Power-law hemolysis threshold parameter sensitivity; Iterative solver convergence tolerance verification; Relative hemolysis index acceptance across five operating points | (none declared) | All seven assessed credibility factors meet or exceed the required gradation levels for MRL 2. Six factors are not assessed, consistent with the ASME V&V 40 MRL 2 assessment matrix for a Class II device with medium model influence. All five flow-speed operating points yield RH < 1.0 at the chosen CFD settings, satisfying the regulatory acceptance criterion. The assessment team concludes that the m… | judge_ready_bundle/packages/adv-2026-p2-115-configuration_low-v04.jsonld | OUT-OF-SCOPE |  | 1 | settings are only in prose |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The packages say it themselves: settings are 'described only in narrative form' and 'no configuration binding node is provided'. The schema has no configuration node to pin values or compare deployment against validation. Maps to Tier-1 candidate W-CX-01. | W-CX-01 (configuration divergence) - direct match
10 | adv-2026-p2-012-w-on-02_low_morrison-cou1-v01 | confirm_existing | W-ON-02 | structural | khakzad/ontological/scope-drift | low | morrison-cou1 | COV-HIT | W-ON-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (W-ON-02 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: CFD solver must be under version control with documented release testing at Gradation 1 or higher. - Numerical code verification [assessed] req1/ach2 AC=yes :: Solver must demonstrate correct numerical implementation via benchmark problems at Gradation 1 or higher. - Discretization error [assessed] req2/ach3 AC=yes :: Mesh convergence study required demonstrating discretization error below 5% for primary QoI (hemolysis index) at Gradation 2. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions must be documented and match intended COU operating conditions (flow rates 2–6 L/min, rotational spee… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted hemolysis index must agree with experimental measurements within 15% across all 5 operating points at Gradation 2. - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: QoIs (hemolysis index, shear stress distribution) must be directly linked to the decision criterion (RH < 1) at Gradation 2. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments must use the same pump geometry, blood analog, and operating conditions as the COU at Gradation 2. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Iterative convergence of residuals to below 1e-5 for all transport equations at Gradation 1. - Use error [not-assessed] req1/ach1 AC=yes :: Evidence of analyst training and simulation setup review at Gradation 1. - Model form [not-assessed] req1/ach1 AC=yes :: Justification for selection of turbulence model and hemolysis power-law model at Gradation 1. - Model inputs [not-assessed] req1/ach1 AC=yes :: Documentation of boundary conditions and material properties with sources at Gradation 1. - Test samples [not-assessed] req1/ach1 AC=yes :: Documentation of blood sample source, preparation, and hematocrit at Gradation 1. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Correspondence between simulation input parameters and experimental test conditions demonstrated at Gradation 1. | CFD hemolysis prediction vs. in-vitro experimental data — FDA generic centrifugal pump; Relative hemolysis index (RH) comparison against predicate device benchmark | (none declared) | All 7 assessed credibility factors meet or exceed required levels for Model Risk Level 2. Six factors are appropriately scoped as not-assessed per ASME V&V 40 Table 4 requirements for MRL 2. CFD hemolysis predictions demonstrate mean agreement of 8.3% against in-vitro benchmarks. The relative hemolysis index criterion (RH < 1) is satisfied across all evaluated operating conditions. The credibility… | judge_ready_bundle/packages/adv-2026-p2-012-w-on-02_low_morrison-cou1-v01.jsonld | CORRECT-DETECTION |  | 1 | hasContextOfUse.operatingEnvelope and applicabilityConstraint both are absent |  | W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. [Ontological] Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
11 | adv-2026-p2-016-w-con-02_high_morrison-cou2-v05 | confirm_existing | W-CON-02 | structural | gohar/requirements/incorrect | high | morrison-cou2 | COV-HIT | W-CON-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-CON-02; W-EP-01; W-EP-02; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (adversarial W-CON-02) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Higher model r… | ProfileMinimal | Accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Discretization error(req4>ach2); Test conditions(req3>ach2); Output comparison(req4>ach2); Relevance of the quantities of interest(req4>ach3); Relevance of the validation activities to the COU(req4>ach2); Numerical solver error(req3>ach1); Use error(req3>ach1); Model form(req4>ach1); Model inputs(req4>ach1); Test samples(req3>ach1); Equivalency of input parameters(req3>ach1) | - Software quality assurance [assessed] req4/ach3 AC=yes :: Version-controlled solver with regression test suite covering hemolysis-relevant flow regimes; formal release process documented. - Numerical code verification [assessed] req4/ach3 AC=yes :: Method of manufactured solutions (MMS) applied to governing equations including non-Newtonian viscosity model; observed order of a… - Discretization error [assessed] req4/ach2 AC=yes :: Richardson extrapolation on three systematically refined meshes; GCI < 5% on hemolysis index QoI; mesh independence demonstrated a… - Test conditions [assessed] req3/ach2 AC=yes :: Experimental boundary conditions match CFD inlet/outlet conditions within ±5%; fluid temperature and hematocrit controlled and doc… - Output comparison [assessed] req4/ach2 AC=yes :: Quantitative comparison metric (e.g., NRMSE < 10%) applied to hemolysis index and velocity field; uncertainty bands from Monte Car… - Relevance of the quantities of interest [assessed] req4/ach3 AC=yes :: QoI (hemolysis index) directly maps to clinical endpoint (plasma free hemoglobin); sensitivity analysis confirms hemolysis index i… - Relevance of the validation activities to the COU [assessed] req4/ach2 AC=yes :: Validation experiments conducted on the same pump geometry at VAD operating conditions (flow 4–6 L/min, speed 2400–3200 RPM); bloo… - Numerical solver error [not-assessed] req3/ach1 AC=yes :: Iterative convergence to residual < 1e-6 on all equations; round-off error analysis performed for hemolysis accumulation path inte… - Use error [not-assessed] req3/ach1 AC=yes :: Independent review of simulation setup by qualified analyst not involved in original simulation; checklist-based verification of b… - Model form [not-assessed] req4/ach1 AC=yes :: Turbulence model selection justified with sensitivity analysis; hemolysis power-law exponents validated against multiple experimen… - Model inputs [not-assessed] req4/ach1 AC=yes :: All material properties and boundary conditions characterized with uncertainty; Monte Carlo input distributions derived from patie… - Test samples [not-assessed] req3/ach1 AC=yes :: Minimum 5 independent experimental replicates per operating point; sample-to-sample variability quantified and propagated to valid… - Equivalency of input parameters [not-assessed] req3/ach1 AC=yes :: Formal equivalency analysis mapping validation experiment conditions to COU conditions; any parameter gaps identified and bounded … | PIV velocity field comparison — centrifugal pump VAD operating point; Monte Carlo hemolysis index UQ validation sweep — 5 flow-speed combinations | (none declared) | Despite multiple credibility factor gaps at MRL 5, the assessment team accepted the package for submission under the condition that all not-assessed factors are resolved prior to regulatory submission. The Monte Carlo UQ results for hemolysis index show predicted values below the clinical threshold at the nominal operating point. Acceptance is conditional and subject to resolution of model form, m… | judge_ready_bundle/packages/adv-2026-p2-016-w-con-02_high_morrison-cou2-v05.jsonld | CORRECT-DETECTION |  | 0.6 | Technically yes the referenceIdentifier does not have a fetch hint but more important there are more load bearing IRIs that are dangling including the requirements in the package, Eiher it should flag all of them or none, not just the optional referenceIdentifier field | This rule needs to be fixed to look at load bearing fields as well | W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. [Consistency] Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
12 | adv-2026-p2-111-inconsistent_medium-v02 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/requirements/inconsistent | medium |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (inconsistent requirements) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: Version-controlled solver with documented test suite; no open critical defects affecting hemolysis QoI. - Numerical code verification [assessed] req1/ach2 AC=yes :: Solver demonstrates correct order of accuracy on canonical benchmark problems relevant to rotating flow. - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for hemolysis index QoI on production mesh; Richardson extrapolation applied across three mesh levels. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions replicated within ±3% of target flow rates; blood analog fluid properties documented. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted RH within 15% of experimentally measured values at all five operating points; systematic bias less than 10%. - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: QoI directly maps to clinical hemolysis endpoint; scalar hemolysis model output linked to relative hemolysis index used in regulat… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation performed on same pump geometry, same operating range, and same fluid as COU; no extrapolation beyond validated envelop… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Residual convergence to 1e-5 for all equations; mass imbalance less than 0.01%. - Use error [not-assessed] req1/ach1 AC=yes :: Simulation setup reviewed by independent analyst; boundary condition specification documented. - Model form [not-assessed] req1/ach1 AC=yes :: Turbulence model selection justified; hemolysis constitutive model documented with literature support. - Model inputs [not-assessed] req1/ach1 AC=yes :: Geometric tolerances, fluid properties, and operating conditions documented with traceability to measurements. - Test samples [not-assessed] req1/ach1 AC=yes :: Minimum three replicate experiments per operating point; sample-to-sample variability documented. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Simulation input parameters demonstrably equivalent to experimental conditions within stated uncertainty bounds. | Bench-scale hemolysis experimental comparison — 5 operating points; Mesh independence study — hemolysis QoI convergence | (none declared) | All seven assessed credibility factors meet or exceed their required levels for MRL 2. Six factors appropriately scoped as not-assessed at MRL 2. Validation results demonstrate RH < 1.0 at all five operating points with mean bias of −4.7% and maximum deviation of 11.4%, within the 15% acceptance threshold. The CFD model is judged credible for identifying worst-case hemolysis operating conditions f… | judge_ready_bundle/packages/adv-2026-p2-111-inconsistent_medium-v02.jsonld | OUT-OF-SCOPE |  | 1 | The conflict in requirements is only expressible in prose and there is no structural way to represent this currently in the context | Refer to INV-19 for more on this | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | B - structural handle exists | PARTIAL: the only package in the gap-probe set carrying hasRequirement nodes (two, each with a specification field). Structure supplies the precondition (two requirements against one claim); the contradiction itself is RH < 1.0 versus RH <= 1.0 inside specification prose. | W-REQ-01 (ambiguous acceptance criterion) - LOOSE FIT: the criteria are contradictory rather than ambiguous
13 | adv-2026-p2-122-theory-preconditions_low-v03 | gap_probe | (none - gap_probe: no rule expected) |  | clarissa-machinery/workflow/theory-preconditions | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: Solver version control and regression testing documented per MRL 2 requirements. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes standard benchmark verification cases relevant to rotating machinery flows. - Discretization error [assessed] req2/ach3 AC=yes :: GCI < 5% for primary quantity of interest (NIH) at representative operating point. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid properties) documented and match CFD boundary conditions within s… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted NIH within ±15% of experimental values at ≥3 of 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI (NIH, RH) directly linked to clinical hemolysis risk metric used in regulatory submission. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry and operating envelope overlap with COU device geometry and CPB operating conditions. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 beyond code verification already performed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; analyst competency assumed via organizational QA. - Model form [assessed] req2/ach2 AC=yes :: Power-law hemolysis model (Giersiepen-Wurzinger) accepted as appropriate for shear-driven hemolysis in centrifugal pumps at MRL 2.… - Model inputs [assessed] req2/ach2 AC=yes :: Material properties (blood viscosity, density) and boundary conditions (inlet flow rate, outlet pressure, rotational speed) docume… - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 given use of published FDA benchmark data. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; input parameter equivalency covered under model inputs assessment. | Mesh convergence study — FDA generic centrifugal pump; Experimental hemolysis index validation against FDA benchmark data; Turbulence model sensitivity — SST k-omega vs. realizable k-epsilon; CFD solver software quality assurance — ANSYS Fluent 2022R2; Lagrangian particle trajectory count sensitivity | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors are appropriately not assessed per ASME V&V 40 Table 3 for MRL 2. Discretization error and output comparison results confirm the CFD model is sufficiently credible to support the regulatory decision that RH < 1 under worst-case CPB operating conditions. The power-law hemolysis model is accepted as the standard… | judge_ready_bundle/packages/adv-2026-p2-122-theory-preconditions_low-v03.jsonld | OUT-OF-SCOPE |  | 1 | Theoretical precondition (linearity of shear-damage accumulation; IID trajectories) stated in description prose only. Distinct from row 12: hasApplicabilityConstraint is a declared property, but ApplicabilityConstraint has zero declared properties, so a precondition still has nowhere structured to go and no way to record whether it was verified for the COU. W-ON-02 fired on the generic absence of any applicability statement, but it fires on 65/71 packages and does not detect this defeater. See INV-19. | Refer to INV-19 for more on this | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The theoretical precondition (linearity and IID across the ensemble) is asserted in description prose. No property records preconditions or whether they were verified for the COU. | NONE - no Tier-1 candidate names this defeater
14 | adv-2026-p2-013-w-si-01_low_nagaraja-cou1-v19 | confirm_existing | W-SI-01 | structural | khakzad/argumentation/undermining-d1 | low | nagaraja-cou1 | GEN-INVALID | W-SI-01 | FALSE |  | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial FEA solver (Ansys) with documented QA procedures; FDA in-house code subject to internal validation protocols. Achieved … - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems from NAFEMS and ASME V&V 10 suite exercised; results within 1% of analytical solutions for linear elastic beam … - Discretization error [assessed] req3/ach3 AC=yes :: Mesh convergence study performed; QoI change <2% between successive refinements. Richardson extrapolation applied to estimate disc… - Numerical solver error [assessed] req2/ach2 AC=yes :: Residual norms monitored; force and displacement residuals converge below 1e-6 relative tolerance. Contact stabilization energy <0… - Use error [assessed] req2/ach2 AC=yes :: Independent analyst review of model setup, boundary conditions, and load application. Checklist-based QA review completed prior to… - Model form [assessed] req3/ach3 AC=yes :: Continuum solid mechanics with nonlinear contact and elasto-plastic titanium alloy material model justified for predicting stiffne… - Model inputs [assessed] req3/ach3 AC=yes :: Material properties (E, yield strength, hardening) sourced from certified material test reports for each lot. Geometric inputs der… - Test samples [assessed] req3/ach3 AC=yes :: N=11 production-representative constructs tested; samples drawn from production lot with documented traceability to device master … - Test conditions [assessed] req3/ach2 AC=yes :: Multiple test conditions (e.g., varied load levels or environmental conditions) required to demonstrate robustness of model predic… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Model input parameters (geometry, material, boundary conditions) confirmed equivalent to physical test conditions through document… - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison metric E/uval < 1.96 and percent difference < 10% for all three QoIs (stiffness, yield force, force at 40 … - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs (stiffness, yield force, force at 40 mm displacement) directly linked to clinical failure modes and regulatory submission req… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario (ASTM F1717 compression-bending, non-cannulated 7.5 mm construct) directly matches the COU geometry, loading m… | Construct stiffness QoI comparison: FEA vs. bench (N=11, two labs); Yield force QoI comparison: FEA vs. bench, E/uval < 1.96 acceptance criterion; Force at 40 mm displacement QoI comparison: FEA vs. bench, <10% difference criterion | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed required levels. The single shortfall (Test conditions, achieved level 2 vs. required level 3) is offset by the direct equivalency of the single ASTM F1717 test configuration to the COU, the multi-laboratory replication across two independent test sites, and the conservative acceptance bounds applied to all QoIs. Overall credibility is s… | judge_ready_bundle/packages/adv-2026-p2-013-w-si-01_low_nagaraja-cou1-v19.jsonld | GENERATOR-ARTIFACT |  | 1 | signature field is missing but the SHACL probably failed the package not allowing the rule to fire |  | W-SI-01 (Medium) — No digital signature. [Structural] Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT. | (no rules fired)
15 | adv-2026-p2-105-model-variance_medium-v05 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/model-variance | medium |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — ML-augmented CPB hemolysis credibility assessment | V&V 40 credibility assessment for a hybrid CFD/ML hemolysis prediction pipeline applied to the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. A convolutional neural network (CNN) surrogate is trained on CFD-generated shear-stress fields to predict the relative hemo… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CNN surrogate training code under version control with documented test suite; CFD solver (Fluent 2022R2) holds FDA-recognized soft… - Numerical code verification [assessed] req2/ach2 AC=yes :: CFD solver order-of-accuracy verified on canonical pipe-flow benchmark; CNN numerical operations verified via gradient-check and u… - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) < 5% for wall shear stress QoI on the CFD training mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental hemolysis measurements conducted under ISO 10993-compliant in-vitro conditions matching CFD boundary conditions withi… - Output comparison [assessed] req2/ach2 AC=yes :: CNN surrogate RH predictions agree with experimental RH measurements within ±10% at all 5 validation operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Model QoI (relative hemolysis index RH) directly maps to the regulatory acceptance criterion and clinical hemolysis risk metric. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments cover the full operating envelope of the COU (flow rates 2–6 L/min, speeds 2500–4000 RPM) and use the same … - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor when discretization error and code verification are assessed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this COU; model is operated by trained analysts following documented SOPs. - Model form [assessed] req2/ach2 AC=yes :: CNN architecture selection justified by comparison of at least two candidate architectures on validation data; selected architectu… - Model inputs [assessed] req2/ach2 AC=yes :: Sensitivity of RH predictions to key input parameters (flow rate, rotational speed, viscosity) characterized; input uncertainty bo… - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor given the in-vitro experimental design. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; boundary condition matching documented under Test Conditions factor. | CNN surrogate point-estimate performance on held-out CFD cases; CNN surrogate vs. experimental hemolysis index — 5 operating points | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. The CNN surrogate achieves MAE=0.031 RH units and R²=0.961 on held-out CFD data, and predictions agree with experimental hemolysis measurements within 7.8% at all 5 COU operating points. Reproducibility of CNN training was reviewed informally and judged adequate by the engineering team. The 6 non-assessed factors are acce… | judge_ready_bundle/packages/adv-2026-p2-105-model-variance_medium-v05.jsonld | OUT-OF-SCOPE |  | 1 | No place to represent this in the context currently |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Seed-to-seed reproducibility is described as 'studied informally and judged adequate' in prose. No property records variance characterization. | NONE - no Tier-1 candidate names this defeater
16 | adv-2026-p2-119-confusion-necessary-sufficient_high-v03 | gap_probe | (none - gap_probe: no rule expected) |  | greenwell/sufficiency/confusion-necessary-sufficient | high |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Version-controlled solver (ANSYS Fluent 2022 R2) with documented release notes and regression test suite; Gradation 1 suffices at … - Numerical code verification [assessed] req1/ach1 AC=yes :: Solver passes vendor-supplied manufactured-solution tests for incompressible Navier-Stokes; Gradation 1 required at MRL 2. - Discretization error [assessed] req2/ach2 AC=yes :: Grid Convergence Index (GCI) for hemolysis scalar QoI below 5% on production mesh; three-level Richardson extrapolation required. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental validation test conditions (flow rate, rotational speed, fluid properties) match COU operating envelope within ±10%. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted relative hemolysis index RH agrees with experimental reference within ±15% across all 5 operating points; systematic bia… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (relative hemolysis index RH) is directly linked to the clinical endpoint; secondary QoIs (shear stress distribution, … - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry (FDA generic pump) is sufficiently similar to COU device geometry; flow regime and fluid properties match COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Iterative convergence residuals below 1e-5 for all transport equations; not required to be formally assessed at MRL 2 beyond basic… - Use error [not-assessed] req1/ach1 AC=yes :: Analyst qualifications and simulation setup review documented; formal use-error assessment not required at MRL 2. - Model form [not-assessed] req1/ach1 AC=yes :: RANS turbulence model selection justified by flow regime; formal model-form uncertainty quantification not required at MRL 2. - Model inputs [not-assessed] req1/ach1 AC=yes :: Nominal input values sourced from device drawings and literature; formal input uncertainty propagation not required at MRL 2. - Test samples [not-assessed] req1/ach1 AC=yes :: Experimental hemolysis data from FDA benchmark study used as reference; sample characterization beyond published metadata not requ… - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Correspondence between simulation inputs and experimental conditions confirmed at a qualitative level; formal equivalency metric n… | Relative hemolysis index benchmark against FDA pump experimental data: RH < 1 satisfied across all 5 operating conditions; Mesh convergence study demonstrating GCI < 2% for shear-stress-derived hemolysis scalar on final mesh; Numerical solver verification against three canonical pipe-flow and rotating-disk benchmarks with residual norms below 1e-6 | (none declared) | The CFD model is accepted as credible for the CPB hemolysis COU on the basis that the relative hemolysis index criterion (RH < 1) is satisfied across all five evaluated operating conditions. The RH criterion is the established regulatory acceptance metric for Class II CPB devices, and its satisfaction by the model — confirmed through validated output comparison at Gradation 2 — constitutes suffici… | judge_ready_bundle/packages/adv-2026-p2-119-confusion-necessary-sufficient_high-v03.jsonld | OUT-OF-SCOPE |  | 1 | This is due to missing constructs to represent AssuranceClaims in Uofa refer to INV-20 for more info |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The sufficiency claim lives in hasDecisionRecord.rationale prose. No property expresses whether a necessary safety property was treated as sufficient for the broader COU goal. | NONE - no Tier-1 candidate names this defeater
17 | adv-2026-p2-120-eliminative-argumentation_medium-v01 | gap_probe | (none - gap_probe: no rule expected) |  | clarissa-machinery/workflow/eliminative-argumentation | medium |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must be under version control with documented test suite covering blood-analog fluid simulations at MRL 2. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation of governing equations through method-of-manufactured-solutions (MMS) or equivalent… - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% at the operating point yielding maximum predicted hemolysis. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions used for validation must be documented with sufficient fidelity to reproduce boundary conditions in t… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted normalised index of hemolysis (NIH) must agree with experimental measurements within ±15% across all five validation ope… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The CFD quantities of interest (hemolysis index, shear stress exposure) must be directly linked to the clinical endpoint (red bloo… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must cover the geometry, fluid properties, and operating range representative of the CPB COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; commercial solver iterative convergence monitoring is considered sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; standard operating procedures and analyst training records are considered sufficient controls. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; selection of turbulence and hemolysis sub-models is documented but formal sensitivity analysis is deferred. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; nominal input values with documented sources are sufficient. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single representative pump specimen used in validation is acceptable. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; qualitative mapping of validation inputs to COU inputs is sufficient. | Bench-scale hemolysis index validation against FDA pump experimental data; Pump pressure-head curve validation across five operating points; Wall shear stress distribution validation against PIV measurements | (none declared) | The credibility assessment for COU1 is accepted based on the totality of positive evidence assembled across the seven assessed factors. Software quality assurance records confirm a controlled, verified solver environment. Numerical code verification against the Taylor-Couette benchmark establishes that the governing equations are correctly implemented. The discretisation error study demonstrates G… | judge_ready_bundle/packages/adv-2026-p2-120-eliminative-argumentation_medium-v01.jsonld | OUT-OF-SCOPE |  | 1 | this defeater needs a declared weakener to attach to, and hasWeakener is [] here and in all 22 gap-probe packages. The rationale argues from 'the totality of positive evidence' only. The structure the defeater needs was never built, refer to INV-21 for more info |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | C - trap 4 (structure never built) | TRAP 4: this defeater needs a declared weakener to attach to, and hasWeakener is [] here and in all 22 gap-probe packages. The rationale argues from 'the totality of positive evidence' only. The structure the defeater needs was never built, so GENERATOR-ARTIFACT is live rather than OUT-OF-SCOPE. | W-AR-06 (eliminative argumentation absent) - direct match
18 | adv-2026-p2-103-coverage-edge-cases_medium-v03 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/coverage-edge-cases | medium |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (synthetic adversarial: coverage gap) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device, including operation at extreme low-flow conditions (≤1 L/min) where recirculation-driven high-shear-stress regions… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Commercial CFD solver (ANSYS Fluent 2022R2) with documented QA procedures and FDA-recognized software validation; version control … - Numerical code verification [assessed] req1/ach1 AC=yes :: Solver passes standard benchmark cases (lid-driven cavity, pipe flow) confirming correct implementation of governing equations. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) < 5% for NIH at nominal operating points; three-mesh Richardson extrapolation performed. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (fluid properties, temperature, inlet boundary conditions) match CFD boundary conditions within ±5% f… - Output comparison [assessed] req2/ach2 AC=yes :: RH < 1.0 at all validated operating points; CFD-predicted NIH within 15% of measured NIH across all nominal test conditions. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (NIH/RH) directly maps to the clinical hemolysis metric; secondary QoI (shear stress distribution) supports mechanisti… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments performed on the FDA generic centrifugal pump geometry under CPB-representative flow and speed conditions; … - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor; commercial solver residual monitoring sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; analyst qualifications and peer review of simulation setup sufficient. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model selection (SST k-omega) justified by literature precedent for rotary blood pumps; hemolysis power-law model appli… - Model inputs [assessed] req2/ach2 AC=yes :: Material properties (blood viscosity, density), boundary conditions (inlet flow rate, rotational speed, outlet pressure), and geom… - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; single pump specimen with documented dimensional inspection sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; direct correspondence between experimental and CFD input parameters documented in simulation plan. | Nominal validation: 2 L/min, 2500 RPM — hemolysis index comparison; Nominal validation: 3 L/min, 2500 RPM — hemolysis index comparison; Nominal validation: 4 L/min, 3000 RPM — hemolysis index comparison; Nominal validation: 5 L/min, 3000 RPM — hemolysis index comparison; Nominal validation: 6 L/min, 3500 RPM — hemolysis index comparison | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors are appropriately not assessed per MRL 2 scoping. Validation evidence across five nominal operating points (2–6 L/min, 2000–3500 RPM) demonstrates RH < 1.0 and NIH prediction error < 15% throughout the nominal CPB clinical envelope. The credibility assessment team determined that the nominal-range validati… | judge_ready_bundle/packages/adv-2026-p2-103-coverage-edge-cases_medium-v03.jsonld | OUT-OF-SCOPE |  | 1 | COU boundary conditions are named in prose. hasContextOfUse carries no operatingEnvelope or applicabilityConstraint, so there is no structured range for a rule to compare validation coverage against. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | COU boundary conditions are named in prose. hasContextOfUse carries no operatingEnvelope or applicabilityConstraint, so there is no structured range for a rule to compare validation coverage against. | NONE - no Tier-1 candidate names this defeater
19 | adv-2026-p2-011-w-on-01_high_nagaraja-cou1-v13 | confirm_existing | W-ON-01 | structural | khakzad/ontological/scope-drift | high | nagaraja-cou1 | GEN-INVALID | W-ON-01 | FALSE |  | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment (Minimal Profile) | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. The FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs across two laboratories. All 13 V&V 40 factors… | ProfileMinimal | [6 empty] | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach3 AC=yes :: Software used in the FEA ensemble must be validated commercial code with documented version control, regression testing, and audit… - Numerical code verification [assessed] req2/ach2 AC=yes :: Code verification benchmarks must demonstrate correct implementation of linear elastic and contact formulations against analytical… - Discretization error [assessed] req3/ach3 AC=yes :: Mesh convergence study must demonstrate that quantities of interest (construct stiffness, peak von Mises stress at screw-rod inter… - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative solver residuals must converge below 1e-6 for all load steps; contact penetration must remain below 1% of minimum elemen… - Use error [assessed] req2/ach2 AC=yes :: Documented simulation workflow with independent analyst review of boundary conditions, material assignments, and load application … - Model form [assessed] req3/ach3 AC=yes :: The constitutive model and contact formulation must be justified by reference to published literature or prior validation studies … - Model inputs [assessed] req3/ach3 AC=yes :: Material properties and geometric inputs must be traceable to measured values or certified material specifications; uncertainty in… - Test samples [assessed] req3/ach3 AC=yes :: Experimental specimens must be production-representative implants with documented traceability to the same design and material lot… - Test conditions [assessed] req3/ach2 AC=yes :: Experimental test conditions must span the range of loading configurations relevant to the COU, including at least two distinct lo… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Model input parameters must be demonstrably equivalent to the experimental specimen parameters; any discrepancies must be quantifi… - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison metric (e.g., validation metric per ASME V&V 20) must be computed; model predictions must fall within expe… - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: The selected QoIs (construct stiffness and yield load) must be directly linked to the clinical performance requirements and regula… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: The validation scenario (F1717 compression-bending, non-cannulated 7.5 mm screw) must match the intended use scenario in geometry,… | F1717 compression-bending stiffness and yield load comparison — FEA vs. experiment (N=11); Cross-laboratory reproducibility assessment for F1717 experimental reference data | (none declared) | The credibility assessment team reviewed all 13 V&V 40 factors. Twelve of thirteen factors meet or exceed the required credibility level for a Model Risk Level 3 submission. The single below-required factor (Test conditions, achieved 2 vs. required 3) reflects the standard single-offset ASTM F1717 protocol and is consistent with accepted regulatory practice for Class II spinal fixation devices. Th… | judge_ready_bundle/packages/adv-2026-p2-011-w-on-01_high_nagaraja-cou1-v13.jsonld | GENERATOR-ARTIFACT |  | 1 | hasContextOfUse missing |  | W-ON-01 (Critical) — No Context of Use at all. [Ontological] Look at: cou_* columns populated? If the package has any COU, firing is wrong. Note: flaw cannot exist in a schema-valid package (COU is sh:minCount mandatory), so a "clean-looking" case targeting this is likely GENERATOR-ARTIFACT (package failed validation upstream). | (no rules fired)
20 | adv-2026-p2-115-configuration_low-v02 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/configuration | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Version-controlled solver with documented release notes; no formal V&V software qualification required at MRL 2. - Numerical code verification [assessed] req1/ach1 AC=yes :: Solver passes standard benchmark cases (lid-driven cavity, pipe flow) confirming correct implementation of governing equations. - Discretization error [assessed] req2/ach2 AC=yes :: GCI between medium and fine mesh below 5% for the primary QoI (normalised hemolysis index). - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid properties) match the five COU operating points within ±5%. - Output comparison [assessed] req2/ach2 AC=yes :: Mean absolute error between CFD-predicted NIH and experimental NIH below 15% across comparison operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (relative hemolysis index RH) is directly linked to the clinical endpoint (blood cell damage) and acceptance criterion… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, operating regime, and fluid properties are representative of the COU device and conditions. - Numerical solver error [assessed] req1/ach1 AC=yes :: Residual convergence below 1e-5 for continuity and momentum; mass imbalance below 0.1%. - Use error [assessed] req1/ach1 AC=yes :: Simulation setup reviewed by a second analyst; boundary conditions and material properties independently verified. - Model form [assessed] req2/ach2 AC=yes :: Turbulence and hemolysis model forms are justified by literature and validated against experimental data for blood pump applicatio… - Model inputs [assessed] req2/ach2 AC=yes :: All model inputs (geometry, fluid properties, boundary conditions, hemolysis model coefficients) are documented and traceable to s… - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor; experimental data sourced from published FDA benchmark dataset. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; boundary condition equivalency addressed under test conditions factor. | Mesh sensitivity study — hemolysis index convergence; Stress-threshold sensitivity — scalar damage accumulation; Temporal sub-step convergence — impeller revolution integration; Experimental comparison — FDA pump hemolysis benchmark | (none declared) | All seven assessed credibility factors meet or exceed the required levels for MRL 2. The six not-assessed factors are acceptable per ASME V&V 40 MRL 2 scoping. CFD-predicted RH values remain below 1.0 across all five COU operating points, supporting the claim that the model correctly identifies worst-case hemolysis conditions. The experimental comparison error of 8.3% is within the 15% acceptance … | judge_ready_bundle/packages/adv-2026-p2-115-configuration_low-v02.jsonld | OUT-OF-SCOPE |  | 1 | The packages say it themselves: settings are 'described only in narrative form' and 'no configuration binding node is provided'. The schema has no configuration node to pin values or compare deployment against validation. Maps to Tier-1 candidate W-CX-01. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The packages say it themselves: settings are 'described only in narrative form' and 'no configuration binding node is provided'. The schema has no configuration node to pin values or compare deployment against validation. Maps to Tier-1 candidate W-CX-01. | W-CX-01 (configuration divergence) - direct match
21 | adv-2026-p2-009-w-al-01_medium_morrison-cou2-v11 | confirm_existing | W-AL-01 | structural | khakzad/aleatoric/randomness | medium | morrison-cou2 | COV-HIT | W-AL-01 | TRUE | COMPOUND-01; W-AL-01; W-AL-02; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Higher model r… | ProfileMinimal | Not accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Test conditions(req5>ach3); Output comparison(req5>ach2); Relevance of the quantities of interest(req5>ach4); Relevance of the validation activities to the COU(req5>ach3); Numerical solver error(req4>ach1); Use error(req4>ach1); Model form(req5>ach1); Model inputs(req5>ach1); Test samples(req4>ach1); Equivalency of input parameters(req5>ach1) | - Software quality assurance [assessed] req4/ach3 AC=yes :: Solver must have documented V&V test suite with regression testing and configuration management at MRL-5 level. - Numerical code verification [assessed] req4/ach3 AC=yes :: Order-of-accuracy tests on canonical flows must confirm expected convergence rates for all governing equations solved. - Discretization error [assessed] req4/ach4 AC=yes :: GCI < 5% for primary QoI on finest mesh; Richardson extrapolation order within 10% of formal order. - Test conditions [assessed] req5/ach3 AC=yes :: Experimental boundary conditions must match COU operating envelope within stated tolerances; blood analog properties must be chara… - Output comparison [assessed] req5/ach2 AC=yes :: Normalized index error < 10% for hemolysis QoI across all operating points; statistical comparison must account for measurement un… - Relevance of the quantities of interest [assessed] req5/ach4 AC=yes :: Primary QoI (NIH hemolysis index) must be directly mapped to clinical endpoint; secondary QoIs must be shown to influence primary … - Relevance of the validation activities to the COU [assessed] req5/ach3 AC=yes :: Validation experiments must be performed on the same device geometry and blood analog under COU-representative operating condition… - Numerical solver error [not-assessed] req4/ach1 AC=yes :: Iterative convergence residuals must be documented and shown to be below solver-specific thresholds for all simulations. - Use error [not-assessed] req4/ach1 AC=yes :: Independent review of simulation setup, boundary condition application, and post-processing scripts must be documented. - Model form [not-assessed] req5/ach1 AC=yes :: Turbulence model and hemolysis damage accumulation model must be justified against literature and sensitivity-tested. - Model inputs [not-assessed] req5/ach1 AC=yes :: All model input parameters must have documented sources and uncertainty characterization; sensitivity ranking must be performed. - Test samples [not-assessed] req4/ach1 AC=yes :: Number of experimental replicates must be sufficient to characterize aleatory variability; statistical power analysis must be docu… - Equivalency of input parameters [not-assessed] req5/ach1 AC=yes :: Mapping between experimental and COU input parameters must be formally demonstrated; any differences must be bounded by sensitivit… | Monte Carlo hemolysis index validation — full operating envelope; Peak wall-shear stress validation at design point; Off-design hemolysis validation — low-flow condition; Pressure-head curve validation across speed range; Shaft torque and hydraulic power validation; Mesh convergence and discretization error assessment | (none declared) | Following review of all 13 credibility factors, the assessment team determines the CFD hemolysis model meets the minimum documentation threshold for submission. While 6 factors remain not-assessed and 7 assessed factors show achieved levels below required levels, the discretization error factor meets its requirement and the Monte Carlo uncertainty analysis provides partial coverage of aleatory var… | judge_ready_bundle/packages/adv-2026-p2-009-w-al-01_medium_morrison-cou2-v11.jsonld | CORRECT-DETECTION |  | 1 | 3 of 6 validation result(s) with no UQ: vad-hemo-peak-shear, vad-hemo-offdesign, vad-torque-power |  | W-AL-01 (High) — ValidationResult with no uncertainty quantification. [Aleatory] Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AL-02 (Medium) — UQ reported but no sensitivity analysis documented. Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
22 | adv-2026-p2-014-w-si-02_low_morrison-cou1-v01 | confirm_existing | W-SI-02 | structural | khakzad/argumentation/undermining-d1 | low | morrison-cou1 | COV-WRONG | W-SI-02 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must be under version control with documented test suite results demonstrating correct implementation of governing equa… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation via manufactured solution or benchmark problem comparison with documented error nor… - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence study with at least three mesh refinement levels; GCI < 5% for hemolysis index QoI. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions must bracket the CFD boundary conditions within ±10% for flow rate and rotational speed. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted hemolysis index must agree with experimental measurements within ±20% across all five operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs computed in simulation must directly correspond to the clinical and regulatory decision metrics for hemolysis assessment. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must use geometry, flow conditions, and fluid properties sufficiently similar to the CPB COU to support extr… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2, iterative convergence monitoring to residuals below 1e-5 is sufficient; formal iterative error quantification not requir… - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2, documented simulation setup review by a second analyst is sufficient; formal use-error analysis not required. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2, qualitative justification of turbulence model and hemolysis correlation selection is sufficient. - Model inputs [not-assessed] req1/ach1 AC=yes :: At MRL 2, nominal input values with documented sources are sufficient; uncertainty propagation not required. - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2, use of published benchmark experimental data with documented provenance is sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2, qualitative mapping between experimental and simulation input parameters is sufficient. | CFD hemolysis index vs. experimental benchmark — FDA generic pump; Relative hemolysis index comparison across 5 flow-speed operating points | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped out per MRL 2 requirements with documented justification. Relative hemolysis index RH < 1.0 confirmed across all five CPB operating points. The CFD model is deemed credible for identifying worst-case hemolysis operating conditions for the Class II CPB device. Assessment accepted by the… | judge_ready_bundle/packages/adv-2026-p2-014-w-si-02_low_morrison-cou1-v01.jsonld | GENERATOR-ARTIFACT |  | 1 | bindsRequirement is actually present in this package why the W-SI-02 rule did not fire |  | W-SI-02 (High) — Missing bindsRequirement. [Structural] Same mandatory-field logic as W-SI-01. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
23 | adv-2026-p2-013-w-si-01_low_morrison-cou1-v03 | confirm_existing | W-SI-01 | structural | khakzad/argumentation/undermining-d1 | low | morrison-cou1 | GEN-INVALID | W-SI-01 | FALSE |  | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Documented version control, regression testing, and release notes for the CFD solver used in hemolysis simulations. - Numerical code verification [assessed] req2/ach2 AC=yes :: Demonstrated solver correctness through benchmark problems with known analytical solutions relevant to blood pump internal flow. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) below 5% for hemolysis index QoI across all five operating points. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid temperature) documented and traceable to simulation boundary cond… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted relative hemolysis index (RH) agrees with experimental measurements within ±15% across all operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The modeled QoI (relative hemolysis index RH) directly maps to the clinical acceptance criterion and regulatory submission metric. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments performed on the same FDA generic centrifugal pump geometry under CPB-representative operating conditions. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor; iterative convergence residuals reported as supplementary information only. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; simulation setup review by a second analyst is sufficient. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; model form selection (k-omega SST turbulence, power-law hemolysis) documented with literature justification… - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; key input parameters (viscosity, density, hemolysis constants) documented with source references. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; sample provenance (bovine blood, certified supplier, lot numbers) documented. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; qualitative argument that bovine blood properties are representative of human blood for hemolysis index com… | CFD hemolysis prediction vs. experimental RH index — 5 operating points; Mesh independence study — three-level refinement for FDA generic pump geometry | (none declared) | All seven assessed credibility factors achieved their required levels under MRL 2. Six factors were appropriately scoped as not-assessed per ASME V&V 40-2018 Table 3 for MRL 2 Class II devices. The CFD model demonstrates adequate credibility to support identification of worst-case hemolysis operating conditions for the FDA generic centrifugal blood pump under CPB conditions. The relative hemolysis… | judge_ready_bundle/packages/adv-2026-p2-013-w-si-01_low_morrison-cou1-v03.jsonld | GENERATOR-ARTIFACT |  | 1 | signatureField is missing but SHACL validation probably failed |  | W-SI-01 (Medium) — No digital signature. [Structural] Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT. | (no rules fired)
24 | adv-2026-p2-114-human-errors_low-v01 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/human-errors | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: Version-controlled CFD solver with documented release notes and regression test suite. - Numerical code verification [assessed] req1/ach2 AC=yes :: Solver verified against at least one canonical benchmark with documented error norms. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) reported for hemolysis index QoI; GCI < 5%. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental flow rates and rotational speeds match simulation boundary conditions within ±2%. - Output comparison [assessed] req2/ach2 AC=yes :: CFD-predicted relative hemolysis index within 15% of experimental measurements at all operating points. - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: QoIs directly map to clinical hemolysis risk metric used in regulatory submission. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry and operating conditions representative of clinical CPB use envelope. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Iterative convergence residuals below 1e-5 for all transport equations. - Use error [not-assessed] req1/ach1 AC=yes :: Not formally defined for this assessment. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model and hemolysis power-law model selection justified with sensitivity analysis. - Model inputs [assessed] req2/ach2 AC=yes :: Blood viscosity, density, and inlet velocity profiles specified from published physiological ranges; sensitivity to input variatio… - Test samples [not-assessed] req1/ach1 AC=yes :: Sample size and variability characterization not required at MRL 2. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Equivalency of simulation input parameters to clinical deployment parameters not required at MRL 2. | Bench hemolysis validation against experimental NIH index data; Pressure-flow characteristic validation; Wall shear stress distribution validation via PIV comparison | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors are not assessed per scoping agreement consistent with MRL 2 requirements. CFD predictions of relative hemolysis index are within 9% of experimental bench measurements across all five operating points, satisfying the RH < 1 acceptance criterion. The evidence package is judged sufficient to support the regu… | judge_ready_bundle/packages/adv-2026-p2-114-human-errors_low-v01.jsonld | OUT-OF-SCOPE |  | 1 | Prose only defeater,cant be reached by rules. description — "The pump is operated by perfusionists during CPB procedures… No operator misuse scenarios, input-entry errors, or training-gap effects are characterized anywhere in the evidence package." validation results — "operator inputs to the test rig assumed to be exact and error-free", "no misuse or off-nominal input conditions were tested or simulated |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | B - structural handle exists | STRUCTURAL: the V&V 40 factor set includes 'Use error', present here with factorStatus=not-assessed, which is a property a rule can read. W-EP-04 covers unassessed factors only when MRL > 2; this is MRL 2, so the engine's silence is per-spec. Do not apply the prose test here - CORRECT-DETECTION is live. | NONE - no Tier-1 candidate names this defeater
25 | adv-2026-p2-118-arguing-from-ignorance_low-v03 | gap_probe | (none - gap_probe: no rule expected) |  | greenwell/sufficiency/arguing-from-ignorance | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Software used for CFD simulation is commercially available with documented version control and quality assurance processes meeting… - Numerical code verification [assessed] req1/ach1 AC=yes :: Numerical code has been verified through standard benchmark problems at Gradation 1. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) computed for the hemolysis index QoI; GCI < 5% on the fine mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions replicate the five target flow-speed operating points within stated tolerances; blood analogue fluid … - Output comparison [assessed] req2/ach2 AC=yes :: Predicted normalized index of hemolysis (NIH) agrees with experimental measurements within 15% at all five operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The hemolysis index QoI is directly linked to the clinical endpoint of red blood cell damage under CPB operating conditions. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments use the same pump geometry and operating range as the COU; blood analogue properties are representative. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Solver residuals monitored; not assessed at MRL 2. - Use error [not-assessed] req1/ach1 AC=yes :: User qualification and simulation execution review; not assessed at MRL 2. - Model form [not-assessed] req1/ach1 AC=yes :: Turbulence model and hemolysis power-law model selection documented; not assessed at MRL 2. - Model inputs [not-assessed] req1/ach1 AC=yes :: Boundary condition uncertainty quantified; not assessed at MRL 2. - Test samples [not-assessed] req1/ach1 AC=yes :: Sample size and variability assessment; not assessed at MRL 2. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Equivalency between model inputs and physical test parameters; not assessed at MRL 2. | CFD hemolysis index comparison against FDA benchmark experimental data; Relative hemolysis index RH < 1 criterion verification across 5 operating points | (none declared) | The CFD hemolysis model for the FDA generic centrifugal blood pump is accepted for use in the CPB COU. Across all seven assessed credibility factors, no failure modes have been observed. No contraindications have been reported in any aspect of the simulation workflow, validation comparisons, or test condition replication. We have not encountered any issues that would undermine confidence in the mo… | judge_ready_bundle/packages/adv-2026-p2-118-arguing-from-ignorance_low-v03.jsonld | OUT-OF-SCOPE |  | 1 | Lack of AssuranceClaims leads to arguing from absence of contradictory evidence. Refer to INV-21 for more info |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Acceptance is argued from absence of contradicting evidence in hasDecisionRecord.rationale ('no failure modes have been observed', 'no contraindications have been reported', 'we have not encountered any issues'). The form of a rationale is not a property. | NONE - no Tier-1 candidate names this defeater
26 | adv-2026-p2-020-w-prov-01_medium_morrison-cou1-v19 | confirm_existing | W-PROV-01 | structural | khakzad/epistemic/incomplete-knowledge | medium | morrison-cou1 | COV-WRONG | W-PROV-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must operate under documented version control and regression testing at Gradation 2. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation of governing equations via benchmark test cases at Gradation 2. - Discretization error [assessed] req2/ach2 AC=yes :: GCI below 5% for primary QoI at Gradation 2. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions must match simulated operating envelope within documented tolerances at Gradation 2. - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric defined and applied; agreement within pre-specified tolerance at Gradation 2. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI must directly map to clinical hemolysis risk indicator at Gradation 2. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation scenario must match COU geometry, fluid, and operating regime at Gradation 2. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2 for this factor; Gradation 1 default accepted. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; Gradation 1 default accepted. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; Gradation 1 default accepted. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; Gradation 1 default accepted. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; Gradation 1 default accepted. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; Gradation 1 default accepted. | Relative hemolysis index CFD vs. experimental benchmark; Mesh independence verification result; Prior experimental hemolysis characterisation summary | (none declared) | All 7 assessed credibility factors meet or exceed required gradation levels for MRL 2. Six factors appropriately scoped out at Gradation 1 per ASME V&V 40 MRL 2 requirements. CFD model is deemed credible for identifying worst-case hemolysis operating conditions for the Class II CPB device. RH index predictions satisfy the RH < 1 acceptance criterion across all 5 operating points. | judge_ready_bundle/packages/adv-2026-p2-020-w-prov-01_medium_morrison-cou1-v19.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 0.7 | The property the rule checks is structurally present and machine-readable; the defeater was planted correctly; and a one-line seed change detects it. That's a rule scoping defect, not a schema gap. |  | W-PROV-01 (Critical) — Chain terminates at a node with no upstream edge and not marked isFoundationalEvidence. [Provenance] Look at: validation_results / evidence chain: any node that is neither foundational-marked nor derived from anything? Scope note from the mutation work: the rule walks upstream FROM claims, so nodes outside any claim's derivation subtree are out of scope; an "orphan" not reachable from a claim doesn't fire, and that is per-spec behavior. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
27 | adv-2026-p2-106-robustness_high-v02 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/robustness | high |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — ML-augmented CPB hemolysis credibility assessment | V&V 40 credibility assessment for a machine-learning surrogate model that predicts pump-induced hemolysis indices for the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. The ML model (a gradient-boosted regressor) was trained on CFD-derived shear-stress histories an… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | (none) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Version-controlled codebase with documented training pipeline and dependency manifest. - Numerical code verification [assessed] req1/ach1 AC=yes :: Underlying CFD solver (ANSYS Fluent 2022 R2) verified against canonical benchmark cases; surrogate training code verified by unit … - Discretization error [assessed] req2/ach2 AC=yes :: Mesh-independence study demonstrating less than 2% change in area-averaged wall shear stress between medium and fine mesh for all … - Test conditions [assessed] req2/ach2 AC=yes :: Experimental validation conditions match COU operating envelope: hematocrit 35±2%, temperature 37±0.5°C, flow rates 2–6 L/min, rot… - Output comparison [assessed] req2/ach2 AC=yes :: Surrogate-predicted RH values agree with experimental RH within ±10% across all nominal validation operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (relative hemolysis index RH) is directly linked to the regulatory acceptance criterion and is the decision-driving ou… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments performed on the same pump platform and operating envelope as the COU; blood analog and test fluid properti… - Numerical solver error [assessed] req1/ach1 AC=yes :: Iterative convergence of CFD residuals to below 1×10⁻⁵ for all transport equations at each operating point. - Use error [assessed] req1/ach1 AC=yes :: Documented user procedures for surrogate inference, including input-range checks that flag extrapolation beyond the training envel… - Model form [assessed] req2/ach2 AC=yes :: Surrogate architecture selection is justified by comparison of at least three candidate model families on the training dataset; se… - Model inputs [assessed] req2/ach2 AC=yes :: Sensitivity analysis identifies the top input drivers of RH prediction; input uncertainty ranges are documented and bounded within… - Test samples [assessed] req1/ach1 AC=yes :: At least three independent pump specimens tested per operating point; specimen variability reported. - Equivalency of input parameters [assessed] req1/ach1 AC=yes :: CFD boundary conditions and experimental test conditions are demonstrated to be equivalent within measurement uncertainty for all … | Held-out test-set RH prediction accuracy; 5-fold cross-validation RMSE on training distribution; Single out-of-distribution geometry evaluation | (none declared) | All 13 ASME V&V 40 credibility factors have been assessed and meet or exceed the required gradation levels for Model Risk Level 2. The gradient-boosted surrogate demonstrates strong predictive accuracy on the held-out test set (MAE = 0.031, R² = 0.961) and five-fold cross-validation (RMSE = 0.038). A single out-of-distribution geometry evaluation returned an absolute error of 0.04, and the assessm… | judge_ready_bundle/packages/adv-2026-p2-106-robustness_high-v02.jsonld | OUT-OF-SCOPE |  | 1 | No property characterizes distribution shift, noise injection or adversarial perturbation. The ValidationResult nodes carry name/description only, so 'not characterized' has nothing to bind to. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | No property characterizes distribution shift, noise injection or adversarial perturbation. The ValidationResult nodes carry name/description only, so 'not characterized' has nothing to bind to. | NONE - no Tier-1 candidate names this defeater
28 | adv-2026-p2-002-w-ar-02_medium_nagaraja-cou1-v18 | confirm_existing | W-AR-02 | D2 | jarzebowicz-wardzinski/argument_defeaters/D2-rebutting | medium | nagaraja-cou1 | COV-HIT | W-AR-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal |  | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req3/ach3 AC=yes :: Validated commercial FEA solver with documented QA procedures and version control; all simulation scripts under configuration mana… - Numerical code verification [assessed] req3/ach3 AC=yes :: Benchmark problems solved with known analytical solutions; residuals and convergence criteria documented. - Discretization error [assessed] req3/ach3 AC=yes :: Mesh convergence study performed; discretization error estimated and shown to be below 5% of QoI. - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative solver convergence verified; round-off error assessed as negligible for double-precision computation. - Use error [assessed] req2/ach2 AC=yes :: Model setup reviewed by independent analyst; boundary conditions and loading confirmed against ASTM F1717 protocol. - Model form [assessed] req3/ach3 AC=yes :: Continuum linear-elastic model form justified for the loading regime; material nonlinearity assessed and shown to be negligible be… - Model inputs [assessed] req3/ach3 AC=yes :: Material properties sourced from certified material test reports; geometry from CAD with dimensional inspection report. - Test samples [assessed] req3/ach3 AC=yes :: N≥10 test specimens from production lots; dimensional measurements confirm conformance to nominal geometry. - Test conditions [assessed] req3/ach2 AC=yes :: Test conditions span the range of physiologically relevant loading scenarios; at minimum two independent load levels or configurat… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Simulation input parameters (geometry, material, boundary conditions) demonstrably equivalent to experimental test configuration. - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison metric defined a priori; E/uval < 1.96 and percent difference < 10% for all primary QoIs. - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs directly correspond to the clinical and regulatory decision quantities; stiffness and yield force are primary design acceptan… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario matches COU in device geometry, loading mode, and regulatory context; no significant extrapolation required. | Construct stiffness QoI comparison — FEA vs. experimental (N=11); Yield force QoI comparison — FEA vs. experimental (N=11); Force at 40 mm displacement QoI comparison — FEA vs. experimental (N=11) | (none declared) | The FEA modeling framework demonstrates adequate credibility for the intended COU. All three primary QoIs satisfy the pre-defined quantitative acceptance criteria (percent difference <10%, E/uval <1.96). Twelve of thirteen V&V 40 credibility factors meet or exceed their required levels. The Test conditions factor achieved Gradation 2 against a required Gradation 3, reflecting the single-condition … | judge_ready_bundle/packages/adv-2026-p2-002-w-ar-02_medium_nagaraja-cou1-v18.jsonld | CORRECT-DETECTION |  | 1 | outcome=Accepted; 1 below required: Test conditions (2<3) |  | W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. [Argumentation] Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
29 | adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v19 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | nagaraja-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Simulation software (Ansys Mechanical 2022 R2) operates under a documented quality management system conforming to ISO 9001. Versi… - Numerical code verification [assessed] req2/ach3 AC=yes :: Gradation 2 requires benchmark comparisons against known analytical solutions for representative load cases. Gradation 3 additiona… - Discretization error [assessed] req3/ach3 AC=yes :: Gradation 3 requires a formal mesh convergence study with at least three refinement levels and a reported Grid Convergence Index (… - Numerical solver error [assessed] req2/ach2 AC=yes :: Gradation 2 requires demonstration that iterative solver residuals converge to specified tolerances and that time-stepping or load… - Use error [assessed] req2/ach2 AC=yes :: Gradation 2 requires a documented modeling protocol reviewed by a second analyst, with evidence that boundary conditions, material… - Model form [assessed] req3/ach3 AC=yes :: Gradation 3 requires documented justification for all constitutive model choices, contact formulations, and geometric simplificati… - Model inputs [assessed] req3/ach3 AC=yes :: Gradation 3 requires that all material and geometric inputs are traceable to measured data from the specific device lot used in be… - Test samples [assessed] req3/ach3 AC=yes :: Gradation 3 requires N≥10 test specimens with documented traceability to production-representative device lots, and statistical ch… - Test conditions [assessed] req3/ach2 AC=yes :: Gradation 3 requires testing across multiple load conditions or configurations representative of the clinical use envelope. Gradat… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Gradation 3 requires quantitative demonstration that model input parameters match experimental conditions within documented uncert… - Output comparison [assessed] req3/ach3 AC=yes :: Gradation 3 requires quantitative comparison metrics (percent difference and validation metric E/uval) computed for each QoI, with… - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: Gradation 2 requires documented traceability from model QoIs to the regulatory or clinical decision being supported. Gradation 3 a… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Gradation 3 requires that the validation scenario matches the COU in device geometry, material, loading mode, and regulatory conte… | Construct stiffness QoI comparison — FEA vs. bench (N=11); Yield force QoI comparison — FEA vs. bench (N=11); Force at 40 mm displacement QoI comparison — FEA vs. bench (N=11) | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed their required levels. The single shortfall (Test conditions, achieved level 2 vs. required level 3) is mitigated by the use of N=11 specimens across two independent laboratories and by regulatory precedent supporting ASTM F1717 single-condition testing for pedicle screw 510(k) submissions. All three QoIs pass quantitative acceptance bou… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v19.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
30 | adv-2026-p2-013-w-si-01_medium_morrison-cou1-v20 | confirm_existing | W-SI-01 | structural | khakzad/argumentation/undermining-d1 | medium | morrison-cou1 | GEN-INVALID | W-SI-01 | FALSE |  | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must be documented under a quality management system with version control and regression testing at Gradation 2. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct numerical implementation via benchmark problems at Gradation 2. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% at Gradation 2. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid properties) must be documented and traceable to calibrated instru… - Output comparison [assessed] req2/ach2 AC=yes :: Normalised hemolysis index predictions must agree with experimental measurements within combined uncertainty bounds at Gradation 2… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs used in validation (NIH, relative hemolysis index RH) must be the same as those used in the COU decision at Gradation 2. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must cover the operating envelope of the COU (flow rates 2–6 L/min, speeds 2000–3500 RPM) at Gradation 2. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal iterative convergence documentation beyond residual monitoring is not required; Gradation 1 is sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2, a formal use-error analysis is not required; Gradation 1 is sufficient. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2, systematic model-form uncertainty analysis is not required; Gradation 1 is sufficient. - Model inputs [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal sensitivity analysis of all model inputs is not required; Gradation 1 is sufficient. - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2, statistical characterisation of test-sample variability is not required; Gradation 1 is sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal equivalency mapping between validation and COU input parameters is not required; Gradation 1 is sufficient. | Five-point bench validation of CFD hemolysis index against experimental measurements on FDA generic centrifugal pump; Mesh convergence study demonstrating discretization error below 3% for shear-stress-derived hemolysis index | (none declared) | All seven assessed credibility factors meet or exceed their required Gradation 2 levels for MRL 2. The six not-assessed factors are acceptable at Gradation 1 per ASME V&V 40 Table 5 for a Class II device at MRL 2. The CFD model is judged credible for identifying worst-case hemolysis operating conditions under CPB use. Note: this package conforms to the Minimal profile; digital signature is pending… | judge_ready_bundle/packages/adv-2026-p2-013-w-si-01_medium_morrison-cou1-v20.jsonld | GENERATOR-ARTIFACT |  | 1 | signatureField is missing but SHACL validation probably failed |  | W-SI-01 (Medium) — No digital signature. [Structural] Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT. | (no rules fired)
31 | adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v08 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | nagaraja-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req2>ach1) | - Software quality assurance [assessed] req2/ach3 AC=yes :: Simulation software (Ansys Mechanical 2022 R2) is under formal configuration management with version-controlled input decks. All s… - Numerical code verification [assessed] req2/ach2 AC=yes :: Gradation 2 requires comparison of solver output against at least one analytical or manufactured solution for a problem class repr… - Discretization error [assessed] req3/ach3 AC=yes :: Gradation 3 requires a systematic mesh convergence study with at least three refinement levels and a Richardson extrapolation-base… - Numerical solver error [assessed] req2/ach2 AC=yes :: Gradation 2 requires monitoring of iterative solver residuals to convergence tolerances specified a priori. Force and displacement… - Use error [assessed] req2/ach2 AC=yes :: Gradation 2 requires a documented simulation plan reviewed by an independent analyst, with a checklist confirming correct boundary… - Model form [assessed] req3/ach3 AC=yes :: Gradation 3 requires documented justification of all constitutive model selections with reference to literature or prior validatio… - Model inputs [assessed] req3/ach3 AC=yes :: Gradation 3 requires that all model inputs are traceable to measured data or published standards, with uncertainty bounds quantifi… - Test samples [assessed] req2/ach3 AC=yes :: Gradation 2 requires that test samples are from the same design family as the simulated device and that sample size is sufficient … - Test conditions [assessed] req2/ach1 AC=yes :: Gradation 2 requires that at least two distinct loading conditions or configurations are tested to bound the model's predictive do… - Equivalency of input parameters [assessed] req2/ach2 AC=yes :: Gradation 2 requires documented mapping of all experimental boundary conditions and loads to their simulation counterparts, with q… - Output comparison [assessed] req3/ach3 AC=yes :: Gradation 3 requires a quantitative uncertainty-aware validation metric (e.g., E/uval) applied to all primary QoIs, with pre-speci… - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: Gradation 2 requires that the QoIs used in validation are the same physical quantities used in the regulatory submission claim. Gr… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Gradation 3 requires that the validation scenario matches the COU in device geometry, material, loading mode, and boundary conditi… | Construct stiffness QoI comparison — FEA vs. bench (N=11); Yield force QoI comparison — FEA vs. bench (N=11); Force at 40 mm displacement QoI comparison — FEA vs. bench (N=11) | (none declared) | The overall credibility assessment for COU1 is Accepted at Medium assurance level. Twelve of thirteen V&V 40 factors meet or exceed their required gradation levels. The single shortfall — Test conditions achieving Gradation 1 against a required Gradation 2 — is mitigated by the direct regulatory applicability of the ASTM F1717 protocol, the strong quantitative validation results across all three Q… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v08.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
32 | adv-2026-p2-106-robustness_low-v02 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/robustness | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — ML hemolysis surrogate credibility assessment (CPB) | V&V 40 credibility assessment for an ML-based surrogate model trained to predict pump-induced hemolysis (relative hemolysis index, RH) for the FDA generic centrifugal blood pump under cardiopulmonary bypass conditions. The surrogate replaces full CFD evaluations for rapid design-space exploration ac… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | (none) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: ML surrogate training pipeline under version control with documented hyperparameter configurations. - Numerical code verification [assessed] req1/ach2 AC=yes :: Surrogate training and inference routines verified against reference implementations on benchmark regression datasets. - Discretization error [assessed] req1/ach1 AC=yes :: CFD-derived training labels generated at mesh density sufficient to achieve grid-converged RH values (GCI < 5%). - Test conditions [assessed] req2/ach2 AC=yes :: Validation test conditions span the five nominal flow-speed operating points specified in the COU. - Output comparison [assessed] req2/ach2 AC=yes :: ML surrogate RH predictions compared against CFD-derived RH labels on test partition; MAE < 0.05 and R² > 0.95. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Surrogate output quantity (RH) is the regulatory-relevant hemolysis metric for CPB device comparison. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities address the COU decision: identifying worst-case hemolysis operating conditions for a Class II CPB device. - Numerical solver error [assessed] req1/ach1 AC=yes :: CFD solver residuals for training-label simulations converge to below 1e-5 for all transport equations. - Use error [assessed] req1/ach1 AC=yes :: Surrogate deployment interface includes input range validation to prevent extrapolation beyond the training envelope. - Model form [assessed] req2/ach2 AC=yes :: Gradient boosting model form is appropriate for the hemolysis response surface given training data characteristics. - Model inputs [assessed] req2/ach2 AC=yes :: Input features are physically motivated, measured with documented uncertainty, and span the nominal COU design space. - Test samples [assessed] req2/ach2 AC=yes :: Test partition is statistically independent from training data and representative of the nominal COU operating space. - Equivalency of input parameters [assessed] req1/ach1 AC=yes :: Input parameters used in surrogate validation match those specified in the COU for the intended CPB operating environment. | Test-set RH prediction accuracy — nominal operating conditions; 5-fold cross-validation RH prediction — nominal dataset; Feature importance and sensitivity analysis — nominal inputs | (none declared) | The ML surrogate for hemolysis prediction satisfies all seven assessed credibility factors at or above required levels for Model Risk Level 2. Test-set performance (MAE=0.031, R²=0.974) meets the acceptance criterion for output comparison. The surrogate correctly identifies worst-case RH operating conditions across the nominal flow-speed design space. All assessed factors are compliant; six factor… | judge_ready_bundle/packages/adv-2026-p2-106-robustness_low-v02.jsonld | OUT-OF-SCOPE |  | 1 | No property characterizes distribution shift, noise injection or adversarial perturbation. The ValidationResult nodes carry name/description only, so 'not characterized' has nothing to bind to. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | No property characterizes distribution shift, noise injection or adversarial perturbation. The ValidationResult nodes carry name/description only, so 'not characterized' has nothing to bind to. | NONE - no Tier-1 candidate names this defeater
33 | adv-2026-p2-004-w-ar-04_low_morrison-cou2-v15 | confirm_existing | W-AR-04 | D4 | jarzebowicz-wardzinski/argument_defeaters/D4-undercutting-model | low | morrison-cou2 | COV-WRONG | W-AR-04 | FALSE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Higher model r… | ProfileMinimal | Not accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach2); Numerical code verification(req4>ach2); Discretization error(req4>ach3); Test conditions(req5>ach3); Output comparison(req5>ach3); Relevance of the quantities of interest(req5>ach4); Relevance of the validation activities to the COU(req5>ach3); Numerical solver error(req4>ach1); Use error(req4>ach1); Model form(req5>ach2); Model inputs(req5>ach2); Test samples(req4>ach1); Equivalency of input parameters(req5>ach1) | - Software quality assurance [assessed] req4/ach2 AC=yes :: Version-controlled solver with documented change log, regression test suite, and FDA 21 CFR Part 11 audit trail for all simulation… - Numerical code verification [assessed] req4/ach2 AC=yes :: Manufactured solution or benchmark convergence study demonstrating formal order of accuracy for hemolysis transport equations. - Discretization error [assessed] req4/ach3 AC=yes :: Richardson extrapolation-based GCI < 2% for NIH QoI on the production mesh at all 5 operating points. - Test conditions [assessed] req5/ach3 AC=yes :: Experimental test conditions (hematocrit, temperature, anticoagulant protocol, pump priming) must match COU operating envelope wit… - Output comparison [assessed] req5/ach3 AC=yes :: Quantitative metric (e.g., E_rms < 10% of clinical threshold) with uncertainty intervals on both simulation and experiment. - Relevance of the quantities of interest [assessed] req5/ach4 AC=yes :: Validated QoI (NIH) must be the same scalar used in the regulatory submission acceptance criterion. - Relevance of the validation activities to the COU [assessed] req5/ach3 AC=yes :: Validation geometry, flow regime, and hemolysis stress exposure history must be representative of the VAD COU operating envelope. - Numerical solver error [not-assessed] req4/ach1 AC=yes :: Iterative convergence residuals < 1e-6 for all transport equations; floating-point round-off analysis documented. - Use error [not-assessed] req4/ach1 AC=yes :: Independent re-run of at least one operating point by a second analyst to detect setup errors in boundary conditions and material … - Model form [assessed] req5/ach2 AC=yes :: Sensitivity study comparing Giersiepen, Taskin, and Fraser hemolysis model forms; justification for selected form documented. - Model inputs [assessed] req5/ach2 AC=yes :: Probability distributions assigned to all uncertain inputs (viscosity, density, hemolysis exponents, inlet velocity profile) with … - Test samples [not-assessed] req4/ach1 AC=yes :: Minimum n=5 independent pump units tested; between-unit variability characterized. - Equivalency of input parameters [not-assessed] req5/ach1 AC=yes :: Formal mapping demonstrating that simulation input parameters (geometry tolerances, surface roughness, fluid properties) are equiv… | Hemolysis index bench validation — FDA pump 5-point operating map; Monte Carlo UQ sweep — hemolysis index uncertainty bands | (none declared) | The credibility assessment team accepted the package for submission review. All 13 credibility factors have been evaluated or flagged. The primary validation evidence (VR-001, VR-002) demonstrates NIH prediction accuracy within acceptable bounds under the 2024-q3-build-2 solver configuration. The team considers the validation results transferable to the current 2026-q1-build-7 model of record, not… | judge_ready_bundle/packages/adv-2026-p2-004-w-ar-04_low_morrison-cou2-v15.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | currentModelVersion='2026-q1-build-7'; validation generated by ['2024-q3-build-2'] — 2 MISMATCH |  | W-AR-04 (High) — Validated at modelVersion ≠ currentModelVersion. [Argumentation] Look at: does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
34 | adv-2026-p2-121-residual-risk-justification_low-v03 | gap_probe | (none - gap_probe: no rule expected) |  | clarissa-machinery/workflow/residual-risk-justification | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Commercial CFD solver (ANSYS Fluent 2022R2) with documented QA procedures and version control. - Numerical code verification [assessed] req1/ach1 AC=yes :: Solver passes standard benchmark cases for rotating flow and scalar transport. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) < 5% for hemolysis index QoI on production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions reproduced in CFD within ±5% for flow rate and rotational speed. - Output comparison [assessed] req2/ach2 AC=yes :: CFD-predicted RHI within ±25% of experimental values at 4 of 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI (RHI) directly maps to clinical hemolysis endpoint used in regulatory submission. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, fluid properties, and operating range span the COU design space. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. | Bench-scale hemolysis index comparison — FDA pump geometry; Mesh independence study — centrifugal pump CFD | (none declared) | The CFD hemolysis model meets all required credibility factor levels for Model Risk Level 2 at the Medium assurance level. Seven factors were formally assessed and all achieved their required levels. Six factors were appropriately scoped out at MRL 2. The model is accepted for use in identifying worst-case hemolysis operating conditions for the Class II CPB device. A known limitation exists in the… | judge_ready_bundle/packages/adv-2026-p2-121-residual-risk-justification_low-v03.jsonld | OUT-OF-SCOPE |  | 1 | limitation is acknowledged but no evaluation of its acceptability for the COU is provided, but hasWeakener is [], so there is no declared weakener for a residual-risk justification to hang off. Refer to INV-21 for more info, this is due to AssuranceClaim not being fleshed out |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | C - trap 4 (structure never built) | TRAP 4: the limitation is acknowledged in description prose ('this limitation is acknowledged but no evaluation of its acceptability for the COU is provided'), but hasWeakener is [], so there is no declared weakener for a residual-risk justification to hang off. GENERATOR-ARTIFACT is live rather than OUT-OF-SCOPE. | W-AR-07 (sustained defeater without residual-risk justification) - direct match
35 | adv-2026-p2-017-w-con-03_low_morrison-cou1-v01 | confirm_existing | W-CON-03 | structural | gohar/requirements/stale | low | morrison-cou1 | COV-WRONG | W-CON-03 | FALSE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (W-CON-03 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented verification test suite and version-controlled source. Release notes and regression test results a… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation of governing equations via manufactured solution or benchmark comparison at require… - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for hemolysis index QoI across all operating points. Minimum three mesh refinement levels required. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (fluid properties, flow rates, rotational speeds) must be documented and match computational boundary… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted RH values must agree with experimental measurements within 20% across all five operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI (relative hemolysis index RH) must be directly linked to the clinical decision criterion and computed from model outputs witho… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments must use the same pump geometry, fluid analog, and operating range as the COU. Geometric fidelity must be c… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2 for this factor; solver residual convergence monitoring is sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; analyst qualification review is sufficient. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; turbulence model selection rationale is sufficient. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; nominal input values with documented sources are sufficient. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single representative pump specimen used in validation is acceptable. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; qualitative equivalency review between validation and COU conditions is sufficient. | Mesh independence study — FDA pump geometry; Relative hemolysis index comparison against experimental data; Turbulence model sensitivity analysis — extended follow-on study | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped to not-assessed at MRL 2. Grid convergence, output comparison, and validation relevance factors provide sufficient confidence in model predictions for the CPB hemolysis COU. The assessment team concludes the CFD model is credible for identifying worst-case hemolysis operating condition… | judge_ready_bundle/packages/adv-2026-p2-017-w-con-03_low_morrison-cou1-v01.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | signatureTimestamp=2024-03-15T09:00:00Z; 3 evidenceTimestamp(s) — 1 AFTER the signature: ['2027-06-22T08:45:00Z'] |  | W-CON-03 (High) — Evidence timestamp later than signature timestamp. [Consistency] Look at: does the package structurally carry BOTH timestamps? Rarely generated (0/180). Absent → GENERATOR-ARTIFACT. Present with evidence > signature → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
36 | adv-2026-p2-104-fidelity_low-v05 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/fidelity | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (fidelity gap variant) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. T… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Commercially validated CFD solver with documented version control and release notes sufficient for MRL 2. - Numerical code verification [assessed] req1/ach1 AC=yes :: Solver passes published benchmark verification tests relevant to internal rotating flow at MRL 2. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) reported for hemolysis-relevant quantities (wall shear stress, RH index) using at least three mesh re… - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid temperature) match COU operating envelope within stated tolerance… - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (e.g., normalized index of agreement or percent error) reported for primary QoI (RH index) across a… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (RH index) is directly linked to the clinical endpoint (hemolysis risk) and the regulatory acceptance criterion (RH < … - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments use the same pump geometry and cover the same operating conditions as the COU. - Numerical solver error [assessed] req1/ach1 AC=yes :: Iterative convergence demonstrated by residual reduction of at least four orders of magnitude for all governing equations. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; qualitative review of analyst competency documentation is sufficient. - Model form [assessed] req2/ach2 AC=yes :: Governing equations and turbulence closure are documented and justified for the flow regime; model form assumptions are stated. - Model inputs [assessed] req2/ach2 AC=yes :: All model boundary conditions and material properties are specified with documented sources and uncertainty ranges. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single experimental dataset from FDA benchmark study is accepted. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; qualitative confirmation that simulation inputs match experimental conditions is sufficient. | Steady-state RH index comparison against FDA benchmark experimental data; Pump pressure head validation at nominal operating conditions | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors are not assessed, which is acceptable per ASME V&V 40 at Model Risk Level 2. The CFD model predicts RH index within 8.3% mean error against FDA benchmark experimental data across all 5 operating points, satisfying the RH < 1 acceptance criterion. The model employs a simplified steady-state Newtonian formulatio… | judge_ready_bundle/packages/adv-2026-p2-104-fidelity_low-v05.jsonld | OUT-OF-SCOPE |  | 1 | The abstraction and its untested impact are stated in prose ('no fidelity trade study documenting the impact of these exclusions has been performed'). No property records abstraction scope or its validation. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The abstraction and its untested impact are stated in prose ('no fidelity trade study documenting the impact of these exclusions has been performed'). No property records abstraction scope or its validation. | NONE - no Tier-1 candidate names this defeater
37 | adv-2026-p2-110-stale_low-v05 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/requirements/stale | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must be under version control with documented release notes and regression test suite passing at MRL 2 per ASME V&V 40-… - Numerical code verification [assessed] req2/ach2 AC=yes :: Code verification must demonstrate correct implementation of governing equations through method of manufactured solutions or equiv… - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence index (GCI) must be computed for primary QoI (hemolysis index) with GCI < 5% at MRL 2 per ASME V&V 40-2018. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions used for validation must be documented and traceable; boundary conditions in simulation must match wi… - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (e.g., E% or validation metric u_val) must be computed for primary QoI at MRL 2 per ASME V&V 40-201… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI used in validation (NIH, RH index) must be directly relevant to the COU decision metric per ASME V&V 40-2018. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must address the same device class, flow regime, and operating envelope as the COU per ASME V&V 40-2018. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2, iterative convergence monitoring to residual < 1e-5 is sufficient; formal solver error quantification not required. - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2, documented simulation procedure with peer review is sufficient; formal use error analysis not required. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2, turbulence model selection must be documented and justified; formal model form uncertainty quantification not required. - Model inputs [not-assessed] req1/ach1 AC=yes :: At MRL 2, nominal input values with stated sources are sufficient; formal input uncertainty propagation not required. - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2, minimum sample characterization (fluid properties, geometry confirmation) is sufficient; statistical sample analysis not… - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2, qualitative equivalency argument between validation and COU input parameters is sufficient. | Relative Hemolysis Index comparison against FDA benchmark data; Pump pressure-head vs. flow-rate curve validation against experimental measurements; Wall shear stress distribution validation against PIV data for FDA generic pump geometry | (none declared) | All 7 assessed credibility factors meet or exceed required gradation levels for MRL 2 per ASME V&V 40-2018. Six factors appropriately scoped out at MRL 2. Discretization error factor exceeds requirement (achieved level 3 vs. required level 2), providing additional confidence. Validation evidence demonstrates CFD model predicts NIH within 12.4% of experimental data across the full COU operating env… | judge_ready_bundle/packages/adv-2026-p2-110-stale_low-v05.jsonld | GENERATOR-ARTIFACT |  | 1 | no manifestation of the defeater at all, neither structural nor prose. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | B - structural handle exists | STRUCTURAL: factorStandard='ASME-VV40-2018' is a machine-readable version string on every factor, so a rule could compare it against a current-edition registry. The staleness assertion itself is in description prose. | W-EV-01 (stale validation data) - LOOSE FIT: the stale item is the standard edition, not the validation data
38 | adv-2026-p2-006-w-ep-02_medium_morrison-cou2-v19 | confirm_existing | W-EP-02 | structural | khakzad/epistemic/incomplete-knowledge | medium | morrison-cou2 | COV-HIT | W-EP-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-02; W-EP-01; W-EP-02; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (W-EP-02 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Monte Carlo un… | ProfileMinimal | Not accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Discretization error(req4>ach3); Test conditions(req4>ach3); Output comparison(req4>ach3); Relevance of the quantities of interest(req5>ach4); Relevance of the validation activities to the COU(req5>ach3); Numerical solver error(req3>ach2); Use error(req3>ach2); Model form(req5>ach2); Model inputs(req4>ach2); Test samples(req4>ach2); Equivalency of input parameters(req4>ach2) | - Software quality assurance [assessed] req4/ach3 AC=yes :: CFD solver must have documented V&V test suite with regression testing and configuration management at level consistent with MRL 5… - Numerical code verification [assessed] req4/ach3 AC=yes :: Formal code verification including method of manufactured solutions or equivalent for all governing equations employed in the hemo… - Discretization error [assessed] req4/ach3 AC=yes :: Grid convergence study with Richardson extrapolation on hemolysis index QoI across at least three systematically refined meshes. - Test conditions [assessed] req4/ach3 AC=yes :: Experimental test conditions must be fully characterized and match COU operating envelope with documented uncertainty in flow rate… - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative comparison metric (e.g., validation metric per ASME V&V 20) applied to hemolysis index predictions against measuremen… - Relevance of the quantities of interest [assessed] req5/ach4 AC=yes :: The QoIs used in validation (hemolysis index) must be identical to those used in the COU decision, with no surrogate quantities su… - Relevance of the validation activities to the COU [assessed] req5/ach3 AC=yes :: Validation experiments must replicate COU geometry, fluid properties, and operating conditions with high fidelity; any deviations … - Numerical solver error [assessed] req3/ach2 AC=yes :: Iterative convergence criteria and floating-point round-off effects must be documented and shown to be negligible relative to disc… - Use error [not-assessed] req3/ach2 AC=yes :: Evidence that the model was applied consistently with developer intent, including documented simulation workflow, boundary conditi… - Model form [not-assessed] req5/ach2 AC=yes :: The hemolysis model form (power-law or mechanistic) must be justified through literature evidence and sensitivity analysis demonst… - Model inputs [not-assessed] req4/ach2 AC=yes :: All model inputs must be characterized with probability distributions and propagated through Monte Carlo UQ; sensitivity analysis … - Test samples [not-assessed] req4/ach2 AC=yes :: Sufficient replicate test samples to characterize inter-sample variability in hemolysis measurements, with statistical power analy… - Equivalency of input parameters [not-assessed] req4/ach2 AC=yes :: Formal equivalency assessment demonstrating that input parameters used in validation experiments match those in the COU, including… | Hemolysis index comparison — operating point 1 (3 L/min, 3000 RPM); Hemolysis index comparison — operating point 2 (4 L/min, 3000 RPM); Hemolysis index comparison — operating point 3 (5 L/min, 3500 RPM); Hemolysis index comparison — operating point 4 (6 L/min, 3500 RPM); Hemolysis index comparison — operating point 5 (4 L/min, 4000 RPM); Velocity field validation against PIV measurements at mid-plane; Pump pressure rise validation across 5 operating points | (none declared) | Despite multiple credibility factor gaps, the assessment team determined that the CFD model provides sufficient evidence to support regulatory submission under a risk-mitigated framework. The Monte Carlo UQ results demonstrate that hemolysis predictions remain below the clinical threshold with 90% confidence at four of five validated operating points. Remaining gaps in model form, test samples, an… | judge_ready_bundle/packages/adv-2026-p2-006-w-ep-02_medium_morrison-cou2-v19.jsonld | CORRECT-DETECTION |  | 1 | 3 of 7 validation result(s) with no wasGeneratedBy |  | W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. [Epistemic] Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
39 | adv-2026-p2-022-compound-03_high_morrison-cou2-v02 | confirm_existing | COMPOUND-03 | structural | khakzad/argumentation/undercutting-d4 | high | morrison-cou2 | COV-HIT | COMPOUND-03 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (COMPOUND-03 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Monte Carlo un… | ProfileMinimal | Accepted | Medium | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | (none) | Discretization error(req4>ach3); Output comparison(req4>ach3); Relevance of the validation activities to the COU(req4>ach3); Use error(req3>ach2); Model form(req5>ach2); Test samples(req3>ach2); Equivalency of input parameters(req4>ach3) | - Software quality assurance [assessed] req3/ach3 AC=yes :: Version-controlled solver with documented regression test suite; all tests passing for the hemolysis-enabled build. - Numerical code verification [assessed] req3/ach3 AC=yes :: Method of manufactured solutions (MMS) applied to Navier-Stokes and hemolysis transport equations; observed order of accuracy with… - Discretization error [assessed] req4/ach3 AC=yes :: Richardson extrapolation on three systematically refined meshes; GCI < 5% for hemolysis index QoI at all operating points. - Test conditions [assessed] req4/ach4 AC=yes :: Experimental boundary conditions (inlet flow rate, outlet pressure, fluid temperature, hematocrit) match COU operating envelope wi… - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative comparison metric (E_rms or similar) computed with uncertainty bounds for all QoIs at all COU operating points; metri… - Relevance of the quantities of interest [assessed] req4/ach4 AC=yes :: Validation QoIs are the same scalar quantities (NIH, volume-averaged shear stress) used in the regulatory decision; no surrogate Q… - Relevance of the validation activities to the COU [assessed] req4/ach3 AC=yes :: Validation experiments performed on the same pump geometry and fluid under COU-representative operating conditions; no geometric o… - Numerical solver error [assessed] req3/ach3 AC=yes :: Iterative convergence demonstrated to at least 10^-5 for all residuals; double-precision arithmetic confirmed. - Use error [assessed] req3/ach2 AC=yes :: Independent review of simulation setup by a second qualified analyst; documented checklist of boundary condition and material prop… - Model form [assessed] req5/ach2 AC=yes :: At least two alternative hemolysis model forms evaluated; model-form uncertainty bounded and incorporated into UQ; sensitivity of … - Model inputs [assessed] req4/ach4 AC=yes :: All material properties and boundary condition inputs traceable to primary sources with stated uncertainty; no assumed values with… - Test samples [assessed] req3/ach2 AC=yes :: Minimum three independent experimental replicates per operating point; inter-replicate variability characterized and reported. - Equivalency of input parameters [assessed] req4/ach3 AC=yes :: Demonstration that validation experiment input parameters are equivalent to COU input parameters within stated tolerances for all … | Flow field validation against FDA PIV benchmark — five operating points; Hemolysis index validation against in-vitro bench measurements; Monte Carlo UQ propagation report — 1000-sample Latin hypercube | (none declared) | The credibility assessment team reviewed the full thirteen-factor evidence package and concluded that the submission provides Medium assurance for the intended regulatory use. The Monte Carlo UQ report demonstrates that 95th-percentile NIH predictions remain below the absolute clinical threshold across all five operating points, providing a quantitative safety margin. The flow-field validation aga… | judge_ready_bundle/packages/adv-2026-p2-022-compound-03_high_morrison-cou2-v02.jsonld | CORRECT-DETECTION |  | 1 | assuranceLevel='Medium' (rule needs != 'Low'); Critical weakeners in the recorded firing list: ['W-AR-02', 'W-EP-01'] |  | COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. [Compound] Look at: assurance_level column ≠ Low AND any Critical in rules_fired. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
40 | adv-2026-p2-016-w-con-02_high_morrison-cou2-v14 | confirm_existing | W-CON-02 | structural | gohar/requirements/incorrect | high | morrison-cou2 | COV-HIT | W-CON-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-05; W-CON-02; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (W-CON-02 adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. All 13 credibi… | ProfileMinimal | Accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | (none) | (none) | - Software quality assurance [assessed] req4/ach4 AC=yes :: CFD solver must have documented V&V test suite, version control, and regression testing consistent with MRL 5 requirements. - Numerical code verification [assessed] req4/ach4 AC=yes :: Order-of-accuracy tests must confirm theoretical convergence rates for momentum and scalar transport equations used in hemolysis p… - Discretization error [assessed] req4/ach4 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% at the operating point of highest hemolysis sensitivity. - Test conditions [assessed] req3/ach3 AC=yes :: Experimental test conditions (fluid properties, pump speed, flow rate, temperature) must be characterized with documented uncertai… - Output comparison [assessed] req4/ach4 AC=yes :: Quantitative metric (e.g., E_rel or normalized index of agreement) for hemolysis index prediction must demonstrate agreement withi… - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: The QoI (normalized index of hemolysis) must be directly mappable to the clinical decision threshold without intermediate surrogat… - Relevance of the validation activities to the COU [assessed] req4/ach4 AC=yes :: Validation experiments must use the same pump geometry, comparable blood analogue properties, and operating conditions within the … - Numerical solver error [assessed] req3/ach3 AC=yes :: Iterative convergence must be demonstrated by reduction of normalized residuals to below 1e-5 for all equations, with monitored Qo… - Use error [assessed] req3/ach3 AC=yes :: Simulation setup must be reviewed by an independent analyst; boundary condition application, turbulence model selection, and hemol… - Model form [assessed] req4/ach4 AC=yes :: The hemolysis power-law model form must be justified by peer-reviewed literature, and sensitivity to model form assumptions (expon… - Model inputs [assessed] req4/ach4 AC=yes :: All model inputs (fluid viscosity, density, RBC fragility coefficients, pump geometry tolerances) must have documented sources wit… - Test samples [assessed] req3/ach3 AC=yes :: Experimental hemolysis measurements must use a minimum of 3 independent blood samples per operating point with documented inter-sa… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Differences between validation experiment input parameters and COU input parameters must be identified and their effect on predict… | PIV flow field validation against FDA benchmark experimental data; Hemolysis index prediction vs. in-vitro measurement across 5 operating points; Monte Carlo uncertainty quantification sensitivity analysis on hemolysis power-law exponents | (none declared) | All 13 credibility factors have been assessed and achieved levels meet or exceed required levels for Model Risk Level 5. Monte Carlo uncertainty quantification confirms that the 95th percentile NIH prediction remains below the clinical acceptance threshold across all COU operating points. The hemolysis power-law methodology is grounded in peer-reviewed literature (see referencesIdentifier). Indepe… | judge_ready_bundle/packages/adv-2026-p2-016-w-con-02_high_morrison-cou2-v14.jsonld | CORRECT-DETECTION |  | 1 | referencesIdentifier: j.ymeth.2024.03.003[DANGLES] |  | W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. [Consistency] Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
41 | adv-2026-p2-005-w-ep-01_low_morrison-cou2-v19 | confirm_existing | W-EP-01 | structural | khakzad/epistemic/incomplete-knowledge | low | morrison-cou2 | COV-HIT | W-EP-01 | TRUE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Higher model r… | ProfileMinimal | Not accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach2); Numerical code verification(req4>ach2); Discretization error(req4>ach3); Test conditions(req5>ach3); Output comparison(req5>ach2); Relevance of the quantities of interest(req4>ach3); Relevance of the validation activities to the COU(req5>ach2); Numerical solver error(req3>ach1); Use error(req3>ach1); Model form(req5>ach1); Model inputs(req4>ach1); Test samples(req4>ach1); Equivalency of input parameters(req4>ach1) | - Software quality assurance [assessed] req4/ach2 AC=yes :: Software quality plan documented with version control, unit testing, and regression testing covering hemolysis model routines. - Numerical code verification [assessed] req4/ach2 AC=yes :: Order-of-accuracy tests and manufactured solution benchmarks demonstrate correct numerical implementation of governing equations. - Discretization error [assessed] req4/ach3 AC=yes :: Grid convergence index (GCI) reported for all QoIs at all five operating points with GCI < 5%. - Test conditions [assessed] req5/ach3 AC=yes :: Experimental test conditions (fluid viscosity, hematocrit, temperature, pump speed, flow rate) match COU operating envelope within… - Output comparison [assessed] req5/ach2 AC=yes :: Quantitative agreement metric (e.g., E_rms) reported with uncertainty bounds for hemolysis index at all five operating points; agr… - Relevance of the quantities of interest [assessed] req4/ach3 AC=yes :: QoIs used in validation (NIH, wall shear stress, residence time) are directly linked to clinical hemolysis threshold decision. - Relevance of the validation activities to the COU [assessed] req5/ach2 AC=yes :: Validation activities performed on geometry, operating conditions, and fluid properties directly representative of the COU VAD dev… - Numerical solver error [not-assessed] req3/ach1 AC=yes :: Iterative convergence criteria documented and demonstrated sufficient for hemolysis QoI; round-off and floating-point error charac… - Use error [not-assessed] req3/ach1 AC=yes :: Documented review process ensures correct model setup, boundary condition specification, and post-processing for each operating po… - Model form [not-assessed] req5/ach1 AC=yes :: Model form uncertainty quantified for turbulence closure and hemolysis damage model; sensitivity to model form choices demonstrate… - Model inputs [not-assessed] req4/ach1 AC=yes :: Monte Carlo UQ on model inputs (viscosity, density, damage model constants) propagated through to hemolysis index with reported co… - Test samples [not-assessed] req4/ach1 AC=yes :: Sufficient replicate experimental measurements available to characterize inter-sample variability in hemolysis index across operat… - Equivalency of input parameters [not-assessed] req4/ach1 AC=yes :: Input parameters used in validation simulations demonstrated equivalent to those of the COU device operating conditions within qua… | PIV velocity field comparison for FDA centrifugal pump; Hemolysis index prediction vs. experimental measurements | (none declared) | Despite multiple credibility factor gaps identified at MRL 5, the assessment team accepted the package on the basis that the CFD model provides qualitative directional accuracy for hemolysis prediction and that supplementary bench testing will be conducted prior to submission. Evidence from PIV validation and hemolysis benchmark comparisons is referenced in the supporting documentation; the bindin… | judge_ready_bundle/packages/adv-2026-p2-005-w-ep-01_low_morrison-cou2-v19.jsonld | CORRECT-DETECTION |  | 1 | claim /nodes[], type='Claim', wasDerivedFrom=ABSENT |  | W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. [Epistemic] Look at: whether the package's claims link to evidence at all. SPECIAL: this rule's guard reads a class (uofa:Claim) the schema never declares, so it CANNOT fire on a schema-conformant package. If target_rule_fired = yes on this rule, the package types its claims non-conformantly → lean GENERATOR-ARTIFACT and note it. If silent on a package whose claims genuinely lack derivation, note "unfireable as shipped" in author_notes; verdict follows Q1 honestly. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
42 | adv-2026-p2-007-w-ep-03_low_morrison-cou2-v17 | confirm_existing | W-EP-03 | structural | khakzad/epistemic/data-imprecision | low | morrison-cou2 | COV-WRONG | W-EP-03 | FALSE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (W-EP-03 stale input data) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. The boundary-c… | ProfileMinimal | Accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Equivalency of input parameters(not-assessed) | Software quality assurance(req3>ach2); Numerical code verification(req3>ach2); Test conditions(req4>ach3); Output comparison(req4>ach3); Relevance of the validation activities to the COU(req4>ach3); Numerical solver error(req3>ach2); Use error(req3>ach2); Model form(req4>ach2); Model inputs(req4>ach3); Test samples(req3>ach2); Equivalency of input parameters(req3>ach2) | - Software quality assurance [assessed] req3/ach2 AC=yes :: Version-controlled solver with documented test suite covering all active physics modules; change-log traceability to validation ba… - Numerical code verification [assessed] req3/ach2 AC=yes :: Order-of-accuracy verification via method of manufactured solutions (MMS) for all governing equations including hemolysis transpor… - Discretization error [assessed] req3/ach3 AC=yes :: Grid convergence index (GCI) below 5% on hemolysis index QoI across three systematically refined meshes. - Test conditions [assessed] req4/ach3 AC=yes :: Experimental test conditions match COU operating envelope within stated measurement uncertainty; temperature, hematocrit, and anti… - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative comparison metric (e.g., E-metric or validation uncertainty) computed with experimental uncertainty propagated; compa… - Relevance of the quantities of interest [assessed] req4/ach4 AC=yes :: Validation QoI is the same scalar (NIH) used as the clinical acceptance criterion; no surrogate metrics substituted. - Relevance of the validation activities to the COU [assessed] req4/ach3 AC=yes :: Validation experiments performed on the same device geometry and fluid at COU operating conditions; no geometric surrogates used. - Numerical solver error [assessed] req3/ach2 AC=yes :: Residual convergence criteria tightened to 1e-6 for all transport equations; iterative convergence demonstrated on QoI. - Use error [assessed] req3/ach2 AC=yes :: Documented simulation execution procedure with independent review of boundary condition setup and post-processing scripts; traceab… - Model form [assessed] req4/ach2 AC=yes :: Sensitivity analysis of hemolysis model exponents (alpha, beta) performed; alternative hemolysis model forms evaluated; model form… - Model inputs [assessed] req4/ach3 AC=yes :: All model input parameters characterized with uncertainty distributions; Monte Carlo propagation performed on NIH QoI; input sensi… - Test samples [assessed] req3/ach2 AC=yes :: Minimum three independent test specimens per operating condition; specimen-to-specimen variability characterized and reported. - Equivalency of input parameters [not-assessed] req3/ach2 AC=yes :: Formal equivalency assessment demonstrating that validation experiment input parameters are within specified tolerance of COU inpu… | Hemolysis index absolute value comparison — 5 operating points; PIV internal velocity field comparison | (none declared) | The credibility assessment team determined that the overall evidence package, while showing gaps in model form uncertainty and several factors below required levels, is sufficient to support the intended regulatory submission under the agreed risk acceptance framework. Monte Carlo UQ results demonstrate NIH predictions remain below clinical threshold with 95% confidence. The team noted that bounda… | judge_ready_bundle/packages/adv-2026-p2-007-w-ep-03_low_morrison-cou2-v17.jsonld | CORRECT-DETECTION |  | 1 | modelRevisionDate=2026-01-10T00:00:00Z; dataVintage(s)=['2018-03-15T00:00:00Z', '2023-06-01T00:00:00Z'] — 2 predate the revision |  | W-EP-03 (High) — Dataset vintage predates model revision date. [Epistemic] Look at: does the package carry BOTH a dataset vintage date AND a model revision date, as dated fields? The chain needs result → activity → used dataset. Most generated packages never build this chain (0/180 measured). If the dates aren't structurally present, the flaw was never instantiated → GENERATOR-ARTIFACT. If both dates present and vintage < revision, firing is correct. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
43 | adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v14 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | nagaraja-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach3 AC=yes :: The simulation software (Ansys Mechanical 2022 R2) must operate under a documented quality management system conforming to ISO 900… - Numerical code verification [assessed] req2/ach2 AC=yes :: Code verification must demonstrate that the solver correctly implements the governing equations for static structural analysis; pa… - Discretization error [assessed] req3/ach3 AC=yes :: A formal mesh-convergence study must demonstrate that QoI values (stiffness, yield force, force at 40 mm) change by less than 2% b… - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative solver residuals must be driven below 1×10⁻⁶ for force and displacement norms; energy balance checks must confirm conser… - Use error [assessed] req2/ach2 AC=yes :: Model setup, boundary condition application, and post-processing scripts must be reviewed by an independent analyst; a documented … - Model form [assessed] req3/ach3 AC=yes :: The constitutive model selection (bilinear isotropic hardening for Ti-6Al-4V) must be justified by literature evidence and confirm… - Model inputs [assessed] req3/ach3 AC=yes :: All material properties and geometric dimensions used in the model must be traceable to measured values from the physical test spe… - Test samples [assessed] req3/ach3 AC=yes :: The validation dataset must comprise at least 10 independent test specimens; specimens must be drawn from production lots with doc… - Test conditions [assessed] req3/ach2 AC=yes :: Validation testing must span at least two distinct loading configurations or boundary condition variants representative of the cli… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Model boundary conditions, loading rates, and fixture geometry must be demonstrated to match the physical test setup within specif… - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison must use a validation metric that accounts for experimental uncertainty; the E/uval metric (ASME V&V 40 Eq… - Relevance of the quantities of interest [assessed] req2/ach3 AC=yes :: The QoIs used in validation (stiffness, yield force, force at 40 mm) must be directly linked to the clinical failure mode of inter… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: The validation activity must be conducted on the same device design and material system as the COU device; surrogate or scaled spe… | Construct stiffness QoI comparison — Laboratory 1 (N=6); Yield force QoI comparison — Laboratory 2 (N=5); Force at 40 mm displacement — pooled N=11 E/uval metric | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed their required gradation levels. The single exception — Test conditions (achieved level 2 against required level 3) — is accepted under the Medium assurance level because the ASTM F1717 compression-bending configuration is the primary regulatory standard for this device class and the model influence is Medium. The overall credibility evi… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v14.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
44 | adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v06 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | nagaraja-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req3/ach3 AC=yes :: SQA plan covers version control, issue tracking, and regression testing for all simulation codes (Ansys Mechanical 2022R2, custom … - Numerical code verification [assessed] req3/ach3 AC=yes :: Code verification achieved when the FEA solver reproduces ASME V&V 10 benchmark solutions for linear elastic beam bending and cont… - Discretization error [assessed] req3/ach3 AC=yes :: Discretization error is acceptable when a formal mesh convergence study demonstrates that QoI values change by less than 2% betwee… - Numerical solver error [assessed] req2/ach2 AC=yes :: Solver error is acceptable when iterative convergence criteria are set to force residual < 1e-4 and displacement residual < 1e-5, … - Use error [assessed] req3/ach3 AC=yes :: Use error is mitigated when a documented simulation workflow with mandatory peer review is followed, and an independent analyst re… - Model form [assessed] req3/ach3 AC=yes :: Model form is adequate when the constitutive model (bilinear isotropic hardening for Ti-6Al-4V) is validated against coupon-level … - Model inputs [assessed] req3/ach3 AC=yes :: Model inputs are adequately characterized when material properties are sourced from lot-specific coupon tests, geometric dimension… - Test samples [assessed] req3/ach3 AC=yes :: Test sample adequacy is established when N >= 10 specimens are tested across at least two independent laboratories, and specimen t… - Test conditions [assessed] req3/ach2 AC=yes :: Test conditions are adequate when at least two independent loading configurations (e.g., compression-bending at two different offs… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Equivalency is established when all model input parameters (geometry, material, boundary conditions, loading) are traceable to exp… - Output comparison [assessed] req3/ach3 AC=yes :: Output comparison is adequate when the validation metric E/uval is computed for all three QoIs using ASME V&V 10.1 methodology and… - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoI relevance is established when each QoI is directly linked to a regulatory submission requirement or a clinically meaningful fa… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation activity relevance is established when the validation scenario (geometry, material, loading, boundary conditions) match… | Construct stiffness QoI comparison — FEA vs. experiment (N=11); Yield force QoI comparison — FEA vs. experiment (N=11); Force at 40 mm displacement QoI comparison — FEA vs. experiment (N=11) | (none declared) | The overall credibility assessment for COU1 is Accepted at Medium assurance level. Twelve of thirteen V&V 40 credibility factors achieved their required level (Gradation 3). One factor (Test conditions) achieved Gradation 2 against a required level of 3, reflecting the single-condition ASTM F1717 test protocol; this shortfall is accepted given regulatory precedent and the direct correspondence bet… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v06.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
45 | adv-2026-p2-015-w-con-01_high_morrison-cou2-v06 | confirm_existing | W-CON-01 | structural | gohar/requirements/inconsistent | high | morrison-cou2 | COV-HIT | W-CON-01 | TRUE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-CON-01; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (Accepted, W-CON-01) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. This package evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. All 13… | ProfileMinimal |  | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | (none) | Numerical code verification(req4>ach3); Test conditions(req4>ach3); Output comparison(req4>ach3); Relevance of the validation activities to the COU(req4>ach3); Model form(req4>ach3); Model inputs(req4>ach3) | - Software quality assurance [assessed] req4/ach4 AC=yes :: Software must be validated under a documented QA plan with traceability to version-controlled solver releases and regression test … - Numerical code verification [assessed] req4/ach3 AC=yes :: Order-of-accuracy verification via method of manufactured solutions (MMS) or equivalent, demonstrating expected convergence rates … - Discretization error [assessed] req4/ach4 AC=yes :: Grid convergence index (GCI) or Richardson extrapolation must demonstrate discretization error below 5% of QoI for all operating p… - Test conditions [assessed] req4/ach3 AC=yes :: Experimental test conditions must be documented with uncertainty bounds and must span the full operating envelope of the COU. - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative comparison metrics (e.g., validation metric, area validation metric) must be computed with uncertainty bands for all … - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: The QoIs used in validation activities must be directly traceable to the clinical decision QoIs defined in the COU. - Relevance of the validation activities to the COU [assessed] req4/ach3 AC=yes :: Validation activities must be performed on geometry, flow regime, and boundary conditions sufficiently representative of the COU d… - Numerical solver error [assessed] req3/ach3 AC=yes :: Iterative convergence must be demonstrated with residuals below solver-recommended thresholds; round-off error must be bounded. - Use error [assessed] req3/ach3 AC=yes :: Simulation setup, boundary condition application, and post-processing must be performed by qualified personnel following documente… - Model form [assessed] req4/ach3 AC=yes :: The mathematical model form, including turbulence closure and hemolysis constitutive law, must be justified by literature evidence… - Model inputs [assessed] req4/ach3 AC=yes :: All model input parameters must be characterized with uncertainty distributions and propagated through the model via formal UQ met… - Test samples [assessed] req3/ach3 AC=yes :: Experimental samples used for validation must be documented with provenance, manufacturing tolerances, and traceability to the val… - Equivalency of input parameters [assessed] reqNone/achNone AC=yes :: Input parameters used in the validation simulations must be demonstrated to be equivalent to those of the COU simulations within a… | PIV velocity field comparison against FDA generic pump experimental data; Hemolysis index benchmark against in-vitro measurements across 5 operating points; Pump pressure-rise validation across operating envelope | (none declared) | Following review of all 13 credibility factors, the assessment team concludes that the CFD hemolysis prediction model is Accepted for the VAD COU. While several factors achieved levels below required levels (Numerical code verification, Test conditions, Output comparison, Relevance of validation activities, Model form, Model inputs), the team determined that the Monte Carlo UQ envelope conservativ… | judge_ready_bundle/packages/adv-2026-p2-015-w-con-01_high_morrison-cou2-v06.jsonld | CORRECT-DETECTION |  | 1 | outcome=Accepted; 1 assessed factor(s) with neither level: Equivalency of input parameters |  | W-CON-01 (High) — Decision Accepted but a factor has neither requiredLevel nor achievedLevel. [Consistency] Look at: decision = Accepted AND any factor with both levels missing (factors_detail). Note the v0.5.14 guard: factors with status 'not-assessed' are exempt; if the only levelless factors are explicitly not-assessed, firing is wrong. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-01 (High) — Decision Accepted but a factor has neither requiredLevel nor achievedLevel. Look at: decision = Accepted AND any factor with both levels missing (factors_detail). Note the v0.5.14 guard: factors with status 'not-assessed' are exempt; if the only levelless factors are explicitly not-assessed, firing is wrong. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
46 | adv-2026-p2-012-w-on-02_high_morrison-cou1-v07 | confirm_existing | W-ON-02 | structural | khakzad/ontological/scope-drift | high | morrison-cou1 | COV-HIT | W-ON-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-CON-04; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileComplete | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented verification test suite and version-controlled release process consistent with MRL 2 requirements. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark solutions for rotating-frame Navier-Stokes and scalar transport (hemolysis power-law model) must match analytical or pub… - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% on the fine-grid solution across all 5 operating points. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions must be documented with measured flow rates, rotational speeds, inlet/outlet pressures, and blood ana… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted relative hemolysis index must fall within ±15% of experimental measurements at all 5 operating points; pressure rise wit… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs used in the model (hemolysis index, shear stress exposure time) must be directly linked to the clinical endpoint (blood traum… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments must use the same pump geometry, fluid properties, and operating range as the COU. Geometric similarity mus… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2 with medium model influence, detailed iterative convergence error quantification beyond residual monitoring is not requir… - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2, a documented simulation workflow with peer review is sufficient; formal use-error FMEA is not required. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2, justification of turbulence and hemolysis model selection is sufficient without formal model-form uncertainty quantifica… - Model inputs [not-assessed] req1/ach1 AC=yes :: At MRL 2, nominal input values with documented sources are sufficient; probabilistic sensitivity analysis is not required. - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2, a single representative pump specimen with documented dimensional inspection is sufficient for validation. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2, qualitative equivalency argument between simulation inputs and experimental conditions is sufficient. | Relative hemolysis index validation against FDA pump experimental data; Pump pressure-rise curve validation across 5 flow-speed operating points; Wall shear stress distribution validation against PIV measurements | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped to level 1 and not assessed per MRL 2 requirements with medium model influence. Discretization error exceeded required level (achieved level 3). Output comparison confirms RH index predictions within acceptance tolerance at all operating points. The computational model is accepted as credi… | judge_ready_bundle/packages/adv-2026-p2-012-w-on-02_high_morrison-cou1-v07.jsonld | CORRECT-DETECTION |  | 1 | COU hasApplicabilityConstraint=ABSENT, hasOperatingEnvelope=ABSENT |  | W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. [Ontological] Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-04 (Medium) — ProfileComplete but no SensitivityAnalysis declared. Look at: profile column = Complete (precondition!) and SA absence. If profile is not Complete, firing is wrong (trap 2 shape). W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
47 | adv-2026-p2-113-faults-software_low-v04 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/faults-software | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | (none) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: The CFD solver and post-processing scripts are under version control with documented release notes; unit tests cover core numerica… - Numerical code verification [assessed] req2/ach2 AC=yes :: Observed order of accuracy within 10% of theoretical order for at least two benchmark problems. - Discretization error [assessed] req2/ach2 AC=yes :: GCI for primary QoI (wall shear stress) less than 5% on the fine mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Bench test flow rates and rotational speeds match the five COU operating points within ±2%. - Output comparison [assessed] req2/ach2 AC=yes :: Mean absolute error in relative hemolysis index RH ≤ 0.15 across all five operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (RH index) directly maps to the clinical hemolysis endpoint; secondary QoIs (shear stress, exposure time) support mech… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry and operating conditions span the COU design space; fluid properties match clinical blood analog. - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative residuals for continuity and momentum reduced by at least four orders of magnitude; mass imbalance less than 0.01%. - Use error [assessed] req1/ach1 AC=yes :: Simulation setup reviewed by a second analyst; boundary conditions and material properties documented and traceable. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model selection justified by comparison with higher-fidelity reference data or literature for rotating pump geometries. - Model inputs [assessed] req2/ach2 AC=yes :: Geometric dimensions, fluid properties, and boundary conditions are traceable to measurement data or published standards with unce… - Test samples [assessed] req1/ach1 AC=yes :: At least one physical pump specimen tested; specimen geometry verified against nominal CAD. - Equivalency of input parameters [assessed] req1/ach1 AC=yes :: Simulation input parameters match bench test conditions within stated tolerances. | Mesh convergence study — nominal completion; Hemolysis bench comparison — nominal completion; Pressure-head characteristic curve validation — nominal completion; Method of manufactured solutions verification — nominal completion | (none declared) | All seven assessed credibility factors meet or exceed their required levels for MRL 2. The six remaining factors are not assessed, which is acceptable at MRL 2 per ASME V&V 40 Table 4. The CFD model demonstrates sufficient credibility to support the regulatory submission decision that the device hemolysis performance meets the RH < 1 acceptance criterion relative to the predicate device across all… | judge_ready_bundle/packages/adv-2026-p2-113-faults-software_low-v04.jsonld | OUT-OF-SCOPE |  | 1 | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. | NONE - no Tier-1 candidate names this defeater
48 | adv-2026-p2-016-w-con-02_low_morrison-cou1-v04 | confirm_existing | W-CON-02 | structural | gohar/requirements/incorrect | low | morrison-cou1 | COV-HIT | W-CON-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-CON-02; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented verification test suite and version-controlled source. Achieved level meets MRL 2 requirement. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation of governing equations via benchmark cases. Achieved level meets MRL 2 requirement. - Discretization error [assessed] req2/ach3 AC=yes :: Grid convergence index (GCI) must be reported and discretization error quantified. Achieved level exceeds MRL 2 requirement. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions must be documented and traceable to simulation boundary conditions. Achieved level meets MRL 2 requir… - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (e.g., normalized RMS error) must be reported for primary QoI. Achieved level meets MRL 2 requireme… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs must be directly linked to the clinical endpoint and regulatory acceptance criterion. Achieved level meets MRL 2 requirement. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must use geometry and operating conditions representative of the COU device and conditions. Achieved level m… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal solver error quantification beyond code verification is not required. - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal use-error analysis is not required. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal model-form uncertainty quantification is not required. - Model inputs [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal input uncertainty propagation is not required. - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2, statistical characterization of test sample variability is not required. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2, formal equivalency mapping between test and simulation inputs is not required. | Relative Hemolysis Index comparison across 5 operating points; Mesh convergence study for FDA generic centrifugal blood pump | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped out per MRL 2 guidance. RHI predictions across all 5 operating conditions satisfy RH < 1 acceptance criterion. CFD model is accepted as credible for identifying worst-case hemolysis operating conditions for the Class II CPB device. | judge_ready_bundle/packages/adv-2026-p2-016-w-con-02_low_morrison-cou1-v04.jsonld | CORRECT-DETECTION |  | 1 | referencesIdentifier: fake-paper[DANGLES] |  | W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. [Consistency] Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
49 | adv-2026-p2-113-faults-software_low-v03 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/faults-software | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (software-intensive CFD pipeline) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. The simulation pipeline is built on a multi-component software stack: a Python-based pre-processing module that generates boundary condition files, ANSY… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial CFD solver with documented QA processes; custom post-processing code under version control with unit tests covering pri… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes standard benchmark cases (lid-driven cavity, Poiseuille flow) with order-of-accuracy consistent with second-order sp… - Discretization error [assessed] req2/ach2 AC=yes :: GCI for primary QoI (RH) below 5% at the selected mesh density. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions (flow rate, rotational speed, fluid viscosity) match CFD inputs within stated measurement uncerta… - Output comparison [assessed] req2/ach2 AC=yes :: RH predictions fall within the 95% prediction interval of experimental data at ≥4 of 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (RH) is directly linked to the regulatory acceptance criterion and clinical hemolysis risk. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments use the same device geometry, fluid, and operating envelope as the COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 beyond code verification activities already performed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 given validation evidence supporting model adequacy. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2. | PIV velocity field comparison — nominal completion; Pump pressure-rise validation — nominal completion; Relative hemolysis index (RH) validation — nominal completion; Mesh sensitivity study — nominal completion | (none declared) | All seven assessed credibility factors met or exceeded the required levels for MRL 2. The CFD hemolysis pipeline — comprising ANSYS Fluent, the C++ Lagrangian post-processor, Python pre-processing, and MATLAB orchestration — produced RH predictions consistent with in-vitro experimental data at four of five operating points, with conservative overprediction at the fifth. Discretization error (GCI 3… | judge_ready_bundle/packages/adv-2026-p2-113-faults-software_low-v03.jsonld | OUT-OF-SCOPE |  | 1 | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. | NONE - no Tier-1 candidate names this defeater
50 | adv-2026-p2-012-w-on-02_medium_morrison-cou1-v06 | confirm_existing | W-ON-02 | structural | khakzad/ontological/scope-drift | medium | morrison-cou1 | COV-HIT | W-ON-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Version-controlled solver with documented QA procedures and regression test suite passing for all relevant modules. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver benchmarked against canonical test cases with documented convergence behavior consistent with stated order of accuracy. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) computed for QoI (hemolysis index) demonstrating discretization error below 5% at the production mesh… - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions (flow rate, rotational speed, fluid properties) match simulation inputs within stated uncertainty… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted normalized index of hemolysis (NIH) within 20% of experimental measurements across all 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI (relative hemolysis index RH and NIH) directly maps to the clinical decision criterion (RH < 1 vs. predicate device). - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments conducted on geometrically similar pump geometry under overlapping operating conditions (flow rate and spee… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Iterative convergence residuals below 1e-5 for all transport equations; no assessment of round-off or iterative error beyond resid… - Use error [not-assessed] req1/ach1 AC=yes :: Analyst competency documented; no independent simulation audit required at MRL 2. - Model form [not-assessed] req1/ach1 AC=yes :: Turbulence model selection documented with literature justification; formal model-form uncertainty quantification not required at … - Model inputs [not-assessed] req1/ach1 AC=yes :: Input parameters (geometry, boundary conditions, fluid properties) documented with sources; sensitivity analysis not required at M… - Test samples [not-assessed] req1/ach1 AC=yes :: Number of experimental replicates documented; statistical characterization of sample variability not required at MRL 2. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Qualitative mapping between validation experiment inputs and COU inputs documented; formal equivalency quantification not required… | CFD hemolysis prediction vs. experimental NIH data — 5 operating points; Relative hemolysis index RH comparison against predicate device across flow-speed matrix | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped to Gradation 1 per MRL 2 requirements and confirmed compliant. CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions for the Class II CPB device. RH < 1 acceptance criterion met across all evaluated operating points. Submission… | judge_ready_bundle/packages/adv-2026-p2-012-w-on-02_medium_morrison-cou1-v06.jsonld | CORRECT-DETECTION |  | 1 | COU hasApplicabilityConstraint=ABSENT, hasOperatingEnvelope=ABSENT |  | W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. [Ontological] Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
51 | adv-2026-p2-014-w-si-02_high_morrison-cou1-v14 | confirm_existing | W-SI-02 | structural | khakzad/argumentation/undermining-d1 | high | morrison-cou1 | COV-WRONG | W-SI-02 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Version-controlled solver with documented QA procedures and regression test suite. - Numerical code verification [assessed] req2/ach2 AC=yes :: Code verified against at least one analytical or manufactured solution relevant to the physics of interest. - Discretization error [assessed] req2/ach3 AC=yes :: Mesh convergence study with GCI < 5% for primary QoI. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions documented with traceability to clinical operating envelope. - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric defined; CFD predictions within acceptable tolerance of experimental measurements. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI directly linked to clinical endpoint (hemolysis) with documented mapping. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities use geometry, fluid properties, and operating conditions representative of the COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor category. | Relative hemolysis index validation against FDA pump experimental data; Mesh convergence study — discretization error quantification; Numerical solver verification against analytical benchmark | (none declared) | All 7 assessed credibility factors meet or exceed required levels for Model Risk Level 2. Six factors appropriately scoped out per MRL 2 requirements for Class II CPB device. CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions. RH < 1 acceptance criterion satisfied across all 5 operating points. Package conforms to ProfileMinimal; re… | judge_ready_bundle/packages/adv-2026-p2-014-w-si-02_high_morrison-cou1-v14.jsonld | GENERATOR-ARTIFACT |  | 1 | bindsRequirement is actually present |  | W-SI-02 (High) — Missing bindsRequirement. [Structural] Same mandatory-field logic as W-SI-01. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
52 | adv-2026-p2-012-w-on-02_medium_nagaraja-cou1-v13 | confirm_existing | W-ON-02 | structural | khakzad/ontological/scope-drift | medium | nagaraja-cou1 | COV-HIT | W-ON-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req2>ach1) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Ansys Mechanical release validation documentation and FDA in-house solver verification records confirm version-controlled, tested … - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems with known analytical solutions used to confirm solver correctness for linear elastic structural mechanics. - Discretization error [assessed] req2/ach2 AC=yes :: Mesh convergence study demonstrates QoI change <2% between successive refinements; Richardson extrapolation applied. - Numerical solver error [assessed] req1/ach1 AC=yes :: Residual norms reported and confirmed below solver default tolerances for all load steps. - Use error [assessed] req2/ach2 AC=yes :: Independent peer review of model setup files and boundary condition application by a second analyst not involved in original model… - Model form [assessed] req3/ach3 AC=yes :: Continuum linear elastic titanium alloy constitutive model with frictional contact at screw-rod interface; model form justified by… - Model inputs [assessed] req3/ach3 AC=yes :: Material properties sourced from ASTM-certified coupon testing of the same Ti-6Al-4V heat; geometric inputs from calibrated CMM me… - Test samples [assessed] req2/ach2 AC=yes :: N=11 production-equivalent constructs from a single manufacturing lot; traceability to device master record confirmed. - Test conditions [assessed] req2/ach1 AC=yes :: At least two distinct load configurations or offset distances tested to characterize model sensitivity to boundary conditions. - Equivalency of input parameters [assessed] req2/ach2 AC=yes :: Simulation inputs traceable to measured specimen properties; no surrogate material substitutions. - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison using E/uval metric with threshold <1.96 and percentage difference <10% for all three QoIs across N=11 spe… - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs (stiffness, yield force, force at 40 mm) directly correspond to regulatory submission endpoints for ASTM F1717 compliance. - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation specimens, loading protocol, and QoIs are directly representative of the intended COU without extrapolation beyond vali… | Construct stiffness QoI comparison — all 11 specimens within 10% bound, E/uval < 1.96; Yield force QoI comparison — FEA ensemble mean within acceptance bounds across both laboratories; Force at 40 mm displacement QoI comparison — pooled E/uval metric satisfies <1.96 threshold | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed their required levels. The single shortfall (Test conditions, achieved level 1 vs. required level 2) is offset by the strength of evidence across all other factors, the direct relevance of the single tested condition to the intended COU, and the conservative acceptance bounds applied to output comparison. The overall credibility evidence… | judge_ready_bundle/packages/adv-2026-p2-012-w-on-02_medium_nagaraja-cou1-v13.jsonld | CORRECT-DETECTION |  | 1 | COU hasApplicabilityConstraint=ABSENT, hasOperatingEnvelope=ABSENT |  | W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. [Ontological] Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
53 | adv-2026-p2-001-w-ar-01_high_morrison-cou1-v17 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | morrison-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: SQA plan documented; version control and regression testing suite in place; solver release notes reviewed for hemolysis-relevant p… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes manufacturer-supplied method-of-manufactured-solutions tests for incompressible Navier-Stokes and scalar transport a… - Discretization error [assessed] req3/ach3 AC=yes :: GCI < 5% on wall shear stress and hemolysis power-law integral between medium and fine mesh; observed order of convergence between… - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions (flow rate, rotational speed, fluid properties) documented with uncertainty bounds; CFD boundary … - Output comparison [assessed] req2/ach3 AC=yes :: Predicted hemolysis index within 20% of experimental measurements at all five operating points; pressure rise within 10% of measur… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs used in validation (NIH, wall shear stress) are the same quantities used in the regulatory submission decision; no surrogate … - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, fluid, and operating range overlap with COU geometry and operating envelope; any gaps documented with enginee… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 when solver is commercially validated and no UDFs are employed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; simulation setup reviewed by second analyst. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model selection justified for rotating machinery hemolysis prediction; power-law hemolysis model applicability document… - Model inputs [assessed] req2/ach2 AC=yes :: All material properties and boundary conditions traceable to measurement data or published sources; sensitivity analysis performed… - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; experimental data sourced from published FDA dataset with documented sample provenance. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 when operating envelope overlap is demonstrated under Relevance of validation activities factor. | Bench-top hemolysis measurement across 5 operating points; Pressure-flow curve validation against FDA pump experimental data; Mesh convergence study demonstrating GCI < 2% on shear stress QoI | (none declared) | All seven assessed credibility factors meet or exceed their required gradation levels at MRL 2. Six factors are appropriately scoped to Gradation 1 and not assessed per MRL 2 guidance. The CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions for the Class II CPB device. The relative hemolysis index criterion (RH < 1) is met across all… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_morrison-cou1-v17.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
54 | adv-2026-p2-119-confusion-necessary-sufficient_medium-v02 | gap_probe | (none - gap_probe: no rule expected) |  | greenwell/sufficiency/confusion-necessary-sufficient | medium |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (adversarial: necessary-sufficient confusion) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal |  |  |  | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: CFD solver is a commercially validated platform with documented version control and regression testing. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems (lid-driven cavity, pipe flow) reproduce analytical solutions within 1% for relevant Reynolds number range. - Discretization error [assessed] req2/ach2 AC=yes :: Richardson extrapolation-based GCI < 5% on shear-stress scalar field across all 5 operating points. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental bench conditions (fluid viscosity, temperature, hematocrit) match CFD boundary conditions within ±10%. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted RH index within ±15% of measured RH index at each of the 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Primary QoI (RH index) is directly linked to the clinical hemolysis endpoint; secondary QoI (shear stress distribution) supports m… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments conducted on the same FDA generic pump geometry at overlapping flow-speed combinations as the COU operating… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; solver residual convergence to 1e-5 documented as proxy. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; analyst training records reviewed as proxy. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; selection of SST k-omega turbulence closure and power-law hemolysis model documented. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; nominal input values from device drawings and literature used. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single lot of bovine blood used for bench experiments. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; dimensional similarity between bench loop and clinical use documented qualitatively. | Bench-scale RH index comparison: CFD vs. experimental hemolysis at 5 operating points; Mesh convergence study for FDA generic centrifugal pump geometry at MRL 2 | (none declared) | The CFD model is accepted for the COU of identifying worst-case hemolysis operating conditions for the Class II CPB device. The assessment team concludes that credibility is established primarily on the basis that the relative hemolysis index (RH) threshold criterion (RH < 1 versus predicate device) is implemented as a quantitative acceptance gate in all simulation outputs. Because the RH threshol… | judge_ready_bundle/packages/adv-2026-p2-119-confusion-necessary-sufficient_medium-v02.jsonld | OUT-OF-SCOPE |  | 1 | No place to represent the AssuranceClaim being made in the decision Record refer to INV-20 for more info |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The sufficiency claim lives in hasDecisionRecord.rationale prose. No property expresses whether a necessary safety property was treated as sufficient for the broader COU goal. | NONE - no Tier-1 candidate names this defeater
55 | adv-2026-p2-005-w-ep-01_low_morrison-cou1-v07 | confirm_existing | W-EP-01 | structural | khakzad/epistemic/incomplete-knowledge | low | morrison-cou1 | COV-HIT | W-EP-01 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial CFD solver with documented verification test suite and version-controlled configuration files. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver benchmarked against at least one canonical flow problem with known analytical solution. - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for hemolysis index QoI across mesh refinement levels. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions documented and traceable; inlet flow rates and rotational speeds within ±3% of CFD boundary condition… - Output comparison [assessed] req2/ach2 AC=yes :: Mean absolute error in RH index ≤ 15% across all validation operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoI (relative hemolysis index RH) is directly linked to the clinical endpoint and regulatory acceptance criterion. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, fluid, and operating conditions are representative of the COU device and flow regime. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. | Mesh independence study — three-level refinement; Relative hemolysis index comparison against FDA experimental data; Turbulence model sensitivity — SST k-omega vs. realizable k-epsilon | (none declared) | All seven assessed credibility factors meet or exceed required levels at Model Risk Level 2. Six factors appropriately scoped out per MRL 2 requirements. Validation evidence demonstrates CFD model adequacy for identifying worst-case hemolysis operating conditions. RH < 1 acceptance criterion satisfied across all five operating points. Model is accepted for CPB use as defined in COU1. | judge_ready_bundle/packages/adv-2026-p2-005-w-ep-01_low_morrison-cou1-v07.jsonld | CORRECT-DETECTION |  | 1 | claim /claims[], type='Claim', wasDerivedFrom=ABSENT |  | W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. [Epistemic] Look at: whether the package's claims link to evidence at all. SPECIAL: this rule's guard reads a class (uofa:Claim) the schema never declares, so it CANNOT fire on a schema-conformant package. If target_rule_fired = yes on this rule, the package types its claims non-conformantly → lean GENERATOR-ARTIFACT and note it. If silent on a package whose claims genuinely lack derivation, note "unfireable as shipped" in author_notes; verdict follows Q1 honestly. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
56 | adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v16 | confirm_existing | W-AR-01 | D1 | jarzebowicz-wardzinski/argument_defeaters/D1-undermining | high | nagaraja-cou1 | COV-WRONG | W-AR-01 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req3/ach3 AC=yes :: Simulation software (Ansys Mechanical 2022 R2) must hold a current ISO 9001 certificate and a documented SQA plan covering version… - Numerical code verification [assessed] req3/ach3 AC=yes :: Each solver in the ensemble must demonstrate second-order spatial convergence on at least three canonical structural benchmarks (p… - Discretization error [assessed] req3/ach3 AC=yes :: Discretization error must be estimated via Richardson extrapolation across at least three mesh densities. The grid convergence ind… - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative solver residuals must converge to below 1e-8 for all load steps. Contact penetration must not exceed 1e-4 mm at any inte… - Use error [assessed] req2/ach2 AC=yes :: All simulation inputs must be reviewed by a second qualified analyst before execution. A documented simulation execution checklist… - Model form [assessed] req3/ach3 AC=yes :: The model must capture nonlinear geometric effects (large-displacement formulation) and frictional contact at all metal-metal inte… - Model inputs [assessed] req3/ach3 AC=yes :: Material properties must be sourced from traceable coupon testing or peer-reviewed literature. Geometric dimensions must be measur… - Test samples [assessed] req3/ach3 AC=yes :: A minimum of N=10 test constructs must be included across at least two independent laboratories. Sample traceability (lot, sterili… - Test conditions [assessed] req3/ach2 AC=yes :: Test conditions must span at least two clinically relevant load magnitudes or offset distances to characterize the model's respons… - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: All simulation boundary conditions must be demonstrably equivalent to the experimental setup, with documented mapping of each expe… - Output comparison [assessed] req3/ach3 AC=yes :: Comparison must use a validation metric that accounts for experimental uncertainty. The E/uval ratio (ASME V&V 10.1) must be below… - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: Each QoI must be directly linked to a clinical failure mode or regulatory acceptance criterion. The linkage must be documented in … - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: The validation scenario must be demonstrated to be representative of the COU with respect to geometry, loading mode, boundary cond… | Construct stiffness QoI comparison — FEA vs. experimental (N=11); Yield force QoI comparison — FEA vs. experimental (N=11); Force at 40 mm displacement QoI comparison — FEA vs. experimental (N=11) | (none declared) | All 13 V&V 40 credibility factors have been assessed. Twelve of thirteen factors achieve the required gradation level. One factor (Test conditions) achieves Gradation 2 against a required Gradation 3, reflecting the single-condition experimental dataset. This shortfall is mitigated by the FEA sensitivity analysis demonstrating bounded QoI variation across the clinical offset range, and by the stre… | judge_ready_bundle/packages/adv-2026-p2-001-w-ar-01_high_nagaraja-cou1-v16.jsonld | GENERATOR-ARTIFACT |  | 1 | 0 factor(s) with a requiredLevel but no acceptanceCriteria — nothing for the rule to catch |  | W-AR-01 (Critical) — Factor has requiredLevel but no acceptance criteria. [Argumentation] Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
57 | adv-2026-p2-016-w-con-02_medium_morrison-cou2-v03 | confirm_existing | W-CON-02 | structural | gohar/requirements/incorrect | medium | morrison-cou2 | COV-HIT | W-CON-02 | TRUE | COMPOUND-01; W-AL-01; W-AR-05; W-CON-02; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (synthetic adversarial) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Monte Carlo un… | ProfileMinimal | Accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | (none) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented version control, regression testing, and release notes demonstrating QA procedures consistent with… - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must pass standard benchmark verification cases (lid-driven cavity, Poiseuille flow) with errors below 1% relative to analy… - Discretization error [assessed] req3/ach3 AC=yes :: Grid convergence index (GCI) for hemolysis index QoI must be below 5% on the fine grid. Richardson extrapolation must be applied a… - Test conditions [assessed] req3/ach3 AC=yes :: Experimental test conditions (flow rate, rotational speed, fluid viscosity, temperature) must be documented with uncertainty bound… - Output comparison [assessed] req3/ach3 AC=yes :: Normalized index of agreement between CFD-predicted and measured hemolysis index must exceed 0.85 across all five operating points… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs used in validation (hemolysis index, shear stress distribution) must be directly traceable to the clinical decision metric (a… - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation experiments must use the same FDA generic pump geometry, blood analog fluid, and operating envelope as the COU. Geometr… - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative convergence must be demonstrated with residuals below 1e-5 for continuity and momentum. Round-off error must be assessed… - Use error [assessed] req2/ach2 AC=yes :: Simulation setup must be independently reviewed by a second qualified analyst. Boundary condition inputs must be traceable to docu… - Model form [assessed] req3/ach3 AC=yes :: Turbulence model selection must be justified via sensitivity analysis comparing at least two turbulence closure models. Power-law … - Model inputs [assessed] req3/ach3 AC=yes :: Uncertainty in key model inputs (viscosity, density, pump geometry tolerances, rotational speed) must be propagated via Monte Carl… - Test samples [assessed] req2/ach2 AC=yes :: At least three independent experimental replicates must be available for each of the five operating points to characterize experim… - Equivalency of input parameters [assessed] req2/ach2 AC=yes :: Dimensional analysis must confirm that non-dimensional groups (Reynolds number, flow coefficient, head coefficient) match between … | FDA generic pump hemolysis index validation against bench measurements; FDA generic pump velocity field validation against PIV data | (none declared) | All 13 credibility factors assessed with achieved levels meeting or exceeding required levels. Monte Carlo uncertainty quantification confirms P90 hemolysis prediction remains below clinical threshold with 18% margin. Independent review completed. Hemolysis model calibration methodology is consistent with established correlation approaches in the rotary blood pump literature. Assessment team accep… | judge_ready_bundle/packages/adv-2026-p2-016-w-con-02_medium_morrison-cou2-v03.jsonld | CORRECT-DETECTION |  | 1 | referencesIdentifier: j.jbiomech.2017.04.021[DANGLES] |  | W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. [Consistency] Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
58 | adv-2026-p2-004-w-ar-04_high_nagaraja-cou1-v08 | confirm_existing | W-AR-04 | D4 | jarzebowicz-wardzinski/argument_defeaters/D4-undercutting-model | high | nagaraja-cou1 | COV-WRONG | W-AR-04 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Ansys Mechanical is a commercially validated solver with documented QA procedures. Internal SOP for FEA execution and review in pl… - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems from ASME V&V 10 suite reproduced within 1% for linear elastic beam bending and contact patch test. - Discretization error [assessed] req3/ach3 AC=yes :: Mesh convergence study demonstrating QoI change < 2% between successive refinement levels. - Numerical solver error [assessed] req2/ach2 AC=yes :: Force residual convergence < 0.5% and energy residual < 0.1% at every load increment. - Use error [assessed] req2/ach2 AC=yes :: Independent peer review of simulation setup by a second qualified analyst; checklist documented. - Model form [assessed] req3/ach3 AC=yes :: Constitutive model captures nonlinear titanium behavior; contact mechanics representation validated against published data. - Model inputs [assessed] req3/ach3 AC=yes :: Material properties sourced from coupon testing of the same lot material; geometry from calibrated CMM measurements. - Test samples [assessed] req3/ach3 AC=yes :: N >= 10 constructs from production lots; dimensional verification performed. - Test conditions [assessed] req3/ach2 AC=yes :: Test conditions span the range of COU loading scenarios; at minimum two load levels evaluated. - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: Simulation inputs (geometry, material, boundary conditions, loading) demonstrably equivalent to test conditions. - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative comparison using E/uval metric with pre-specified acceptance bound; all three QoIs evaluated. - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs directly linked to clinical failure modes and regulatory acceptance criteria for pedicle screw constructs. - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario matches COU device, loading mode, and intended use environment. | Construct stiffness QoI comparison — non-cannulated 7.5 mm pedicle screw; Yield force QoI comparison — non-cannulated 7.5 mm pedicle screw | (none declared) | The credibility assessment for COU1 (non-cannulated 7.5 mm pedicle screw compression-bending) is accepted at Medium assurance level. Twelve of thirteen V&V 40 credibility factors meet or exceed the required gradation level. The single exception — Test conditions (Gradation 2 achieved vs. Gradation 3 required) — is acknowledged as a gap; however, the ASTM F1717 single-condition test is the regulato… | judge_ready_bundle/packages/adv-2026-p2-004-w-ar-04_high_nagaraja-cou1-v08.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | currentModelVersion='ansys-mechanical-2026-q1-build-7'; validation generated by ['ansys-mechanical-2026-q1-build-7-rc2'] — 2 MISMATCH |  | W-AR-04 (High) — Validated at modelVersion ≠ currentModelVersion. [Argumentation] Look at: does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
59 | adv-2026-p2-101-data-drift_low-v05 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/evidence_validity/data-drift | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (expanded envelope) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). … | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach1 AC=yes :: Commercial CFD solver with documented QA procedures and version control meets MRL 2 requirement of Gradation 1. - Numerical code verification [assessed] req1/ach1 AC=yes :: Benchmark test cases from ASME V&V 20 suite demonstrate solver correctness at Gradation 1. - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for hemolysis index QoI across all assessed operating points. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions match simulation boundary conditions within stated uncertainty bands for all five original operating … - Output comparison [assessed] req2/ach2 AC=yes :: CFD-predicted relative hemolysis index agrees with in-vitro measurements within ±15% for all five adult operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: Hemolysis index and shear-stress distribution are the primary QoIs; both are directly linked to the clinical decision criterion RH… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities cover the geometry, fluid properties, and operating conditions representative of the COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; commercial solver residual convergence criteria applied. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; standard analyst training and peer review of simulation setup are applied. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model and hemolysis power-law model are justified by literature and sensitivity analysis at Gradation 2. - Model inputs [assessed] req2/ach2 AC=yes :: Inlet velocity profiles, fluid viscosity, and rotational speed inputs are characterised with stated uncertainty at Gradation 2. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single representative pump geometry used. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; direct parameter mapping from test to simulation documented. | FDA generic pump adult-range hemolysis validation (2021) — flow rates 2.0–5.0 L/min, rotational speeds 2000–3500 RPM; Mesh-independence and discretization error study (2021) | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors are not assessed, consistent with MRL 2 scoping. The assessment team acknowledges that the operating envelope has expanded since the 2021 validation campaign to include paediatric and high-flow adult perfusion conditions; however, the team judges the existing validation evidence sufficient by analogy and g… | judge_ready_bundle/packages/adv-2026-p2-101-data-drift_low-v05.jsonld | OUT-OF-SCOPE |  | 1 | Envelope expansion and missing re-calibration are stated in prose. The package carries no dataset vintage and no model revision date, so W-EP-03's date comparison has nothing to read. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Envelope expansion and missing re-calibration are stated in prose. The package carries no dataset vintage and no model revision date, so W-EP-03's date comparison has nothing to read. | W-EV-01 (stale validation data) - direct match
60 | adv-2026-p2-022-compound-03_medium_morrison-cou1-v06 | confirm_existing | COMPOUND-03 | structural | khakzad/argumentation/undercutting-d4 | medium | morrison-cou1 | COV-HIT | COMPOUND-03 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Discretization error(req2>ach1) | - Software quality assurance [assessed] req2/ach2 AC=yes :: CFD solver must have documented version control, regression testing, and release notes consistent with MRL 2 requirements. - Numerical code verification [assessed] req2/ach2 AC=yes :: Code verification must demonstrate correct implementation of governing equations through at least one manufactured solution or ben… - Discretization error [assessed] req2/ach1 AC=yes :: Grid convergence study must be performed with at least three systematically refined meshes and GCI reported for the hemolysis inde… - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions used for validation must be documented with uncertainty bounds on flow rate, rotational speed, and fl… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted hemolysis index and pressure-head must be compared against experimental data with quantitative agreement metrics reporte… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The QoI used in simulation (relative hemolysis index RH) must directly correspond to the clinical and regulatory acceptance criter… - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation activities must be performed on geometry and operating conditions sufficiently similar to the COU device and intended o… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor; default level 1 acceptable. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; default level 1 acceptable. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; default level 1 acceptable. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; default level 1 acceptable. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; default level 1 acceptable. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; default level 1 acceptable. | Bench-level hemolysis validation against FDA pump experimental data; Pressure-head curve validation across 5 operating points | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped out at MRL 2. Overall credibility assessment supports Medium assurance level for the intended COU. Model is accepted for use in identifying worst-case hemolysis operating conditions for the Class II CPB device. | judge_ready_bundle/packages/adv-2026-p2-022-compound-03_medium_morrison-cou1-v06.jsonld | CORRECT-DETECTION |  | 1 | assuranceLevel='Medium' (rule needs != 'Low'); Critical weakeners in the recorded firing list: ['W-AR-02', 'W-EP-01'] |  | COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. [Compound] Look at: assurance_level column ≠ Low AND any Critical in rules_fired. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
61 | adv-2026-p2-014-w-si-02_low_morrison-cou1-v06 | confirm_existing | W-SI-02 | structural | khakzad/argumentation/undermining-d1 | low | morrison-cou1 | COV-WRONG | W-SI-02 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Software used in the simulation is documented, version-controlled, and validated for cardiovascular CFD applications. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver correctness is confirmed through benchmark cases with known analytical solutions. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) for the hemolysis index quantity of interest is below 5%. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions replicate the five target flow-speed operating points within stated tolerances. - Output comparison [assessed] req2/ach3 AC=yes :: Predicted normalised index of hemolysis (NIH) agrees with experimental measurements within 15% at all five operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The QoI (relative hemolysis index RH) is directly linked to the clinical endpoint of pump-induced haemolysis. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments use the same pump geometry, blood analogue fluid, and operating envelope as the COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; iterative convergence monitoring is sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; standard operating procedures govern simulation execution. - Model form [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; model form selection is justified by literature precedent. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; nominal input values from literature are used. - Test samples [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; single representative test article is sufficient. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not assessed at MRL 2; documented correspondence between simulation and test inputs is sufficient. | Hemolysis index bench-test comparison against FDA pump experimental data; Mesh convergence study for rotating impeller domain | (none declared) | All seven assessed credibility factors meet or exceed their required levels at MRL 2. The six non-assessed factors are appropriately scoped out under MRL 2 requirements. The CFD model is judged credible for identifying worst-case hemolysis operating conditions for the Class II CPB device. RH < 1 criterion is satisfied at all five operating points. | judge_ready_bundle/packages/adv-2026-p2-014-w-si-02_low_morrison-cou1-v06.jsonld | GENERATOR-ARTIFACT |  | 1 | bindsRequirement is not missing |  | W-SI-02 (High) — Missing bindsRequirement. [Structural] Same mandatory-field logic as W-SI-01. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
62 | adv-2026-p2-019-w-con-05_high_morrison-cou1-v11 | confirm_existing | W-CON-05 | structural | gohar/requirements/inconsistent | high | morrison-cou1 | COV-HIT | W-CON-05 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-CON-05; W-EP-01; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Version-controlled solver with documented build and test procedures; regression test suite passing at release. - Numerical code verification [assessed] req2/ach2 AC=yes :: Demonstrated second-order spatial convergence on at least two benchmark problems using the method of manufactured solutions. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) for QoI (pressure head, hemolysis index) below 5% on the production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions documented and replicated in simulation within stated uncertainty bounds. - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (e.g., E% or NRMSE) computed for all QoIs with documented acceptance threshold. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs (RH index, pressure head) are directly linked to the clinical decision criterion and regulatory acceptance threshold. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation experiments conducted at flow conditions, geometric scale, and fluid properties representative of the CPB COU. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. | Pressure head validation at 5 operating points; Relative hemolysis index validation; Velocity field comparison via PIV | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. Validation evidence demonstrates CFD model predictions are within acceptance criteria for pressure head (4.2% error), velocity field (6.8% NRMSE), and relative hemolysis index (within 95% CI). Six factors are appropriately scoped out at MRL 2. The model is credible for identifying worst-case hemolysis operating conditions… | judge_ready_bundle/packages/adv-2026-p2-019-w-con-05_high_morrison-cou1-v11.jsonld | CORRECT-DETECTION |  | 1 | VerificationActivity declared but no Evidence linked via wasGeneratedBy. |  | W-CON-05 (High) — VerificationActivity declared but no Evidence linked via wasGeneratedBy. [Consistency] Look at: any declared verification activity; does evidence point back to it? Activity with nothing generated = procedural placeholder, correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-05 (High) — VerificationActivity declared but no Evidence linked via wasGeneratedBy. Look at: any declared verification activity; does evidence point back to it? Activity with nothing generated = procedural placeholder, correct firing. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
63 | adv-2026-p2-115-configuration_low-v03 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/configuration | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req1/ach2 AC=yes :: CFD solver must operate under documented version control with regression test suite confirming reproducible outputs. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver must demonstrate correct implementation of governing equations via manufactured-solution or benchmark comparison. - Discretization error [assessed] req2/ach2 AC=yes :: GCI for the primary QoI (relative hemolysis index RH) must be less than 5% at the production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions for PIV and hemolysis measurements must be documented and traceable to the computational boundary con… - Output comparison [assessed] req2/ach2 AC=yes :: Quantitative comparison metric (e.g., normalised RMSE) for velocity and hemolysis QoIs must be below acceptance thresholds defined… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The QoIs used in validation (velocity, RH) must directly support the decision criterion (worst-case hemolysis identification). - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, flow regime, and fluid properties must be sufficiently similar to the COU device and operating conditions. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: At MRL 2 a formal solver error budget is not required; awareness of dominant solver error sources is sufficient. - Use error [not-assessed] req1/ach1 AC=yes :: At MRL 2 formal use-error analysis is not required. - Model form [not-assessed] req1/ach1 AC=yes :: At MRL 2 a formal model-form uncertainty quantification is not required; model selection rationale must be documented. - Model inputs [assessed] req2/ach2 AC=yes :: All model input parameters (geometry, boundary conditions, material properties) must be traceable to measurements or published sou… - Test samples [not-assessed] req1/ach1 AC=yes :: At MRL 2 formal test-sample characterisation beyond standard fluid property measurement is not required. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: At MRL 2 a formal equivalency argument for input parameters between validation and deployment is not required. | Mesh convergence study — hemolysis index; PIV velocity field comparison; Hemolysis power-law threshold calibration | (none declared) | All seven assessed credibility factors meet or exceed their required levels for MRL 2. The six not-assessed factors are acceptable at MRL 2 per ASME V&V 40-2018 Table 1. The CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions for the Class II CPB device. Validation results obtained at the chosen configuration settings for mesh elemen… | judge_ready_bundle/packages/adv-2026-p2-115-configuration_low-v03.jsonld | OUT-OF-SCOPE |  | 1 | The packages say it themselves: settings are 'described only in narrative form' and 'no configuration binding node is provided'. The schema has no configuration node to pin values or compare deployment against validation. Maps to Tier-1 candidate W-CX-01. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | The packages say it themselves: settings are 'described only in narrative form' and 'no configuration binding node is provided'. The schema has no configuration node to pin values or compare deployment against validation. Maps to Tier-1 candidate W-CX-01. | W-CX-01 (configuration divergence) - direct match
64 | adv-2026-p2-020-w-prov-01_low_morrison-cou2-v15 | confirm_existing | W-PROV-01 | structural | khakzad/epistemic/incomplete-knowledge | low | morrison-cou2 | COV-WRONG | W-PROV-01 | FALSE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment (W-PROV-01 orphan chain node) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. This package c… | ProfileMinimal |  |  |  | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Discretization error(req4>ach3); Test conditions(req4>ach2); Output comparison(req4>ach2); Relevance of the quantities of interest(req4>ach3); Relevance of the validation activities to the COU(req4>ach2); Numerical solver error(req3>ach1); Use error(req3>ach1); Model form(req4>ach1); Model inputs(req4>ach1); Test samples(req3>ach1); Equivalency of input parameters(req3>ach1) | - Software quality assurance [assessed] req4/ach3 AC=yes :: Software must be under full configuration management with documented V&V test suite covering all solver modules. - Numerical code verification [assessed] req4/ach3 AC=yes :: Order-of-accuracy verification using method of manufactured solutions for all governing equations including hemolysis transport. - Discretization error [assessed] req4/ach3 AC=yes :: Richardson extrapolation-based GCI reported for hemolysis index QOI on at least three systematically refined meshes. - Test conditions [assessed] req4/ach2 AC=yes :: Experimental test conditions must match COU operating envelope within stated uncertainty bounds for all 5 flow-speed combinations. - Output comparison [assessed] req4/ach2 AC=yes :: Quantitative metric (e.g., E_rel) reported with uncertainty bands; prediction must lie within experimental uncertainty for all QOI… - Relevance of the quantities of interest [assessed] req4/ach3 AC=yes :: QOIs directly map to absolute hemolysis index values used in the regulatory decision without intermediate surrogate transformation… - Relevance of the validation activities to the COU [assessed] req4/ach2 AC=yes :: Validation activities performed on geometry, fluid, and operating conditions representative of the target VAD COU. - Numerical solver error [not-assessed] req3/ach1 AC=yes :: Iterative convergence criteria and floating-point round-off effects documented and bounded for hemolysis QOI. - Use error [not-assessed] req3/ach1 AC=yes :: Documented review process ensuring correct model setup, boundary condition application, and post-processing for all simulation run… - Model form [not-assessed] req4/ach1 AC=yes :: Justification of turbulence model and hemolysis constitutive model selection with sensitivity analysis demonstrating low model-for… - Model inputs [not-assessed] req4/ach1 AC=yes :: All uncertain model input parameters identified, characterized with probability distributions, and propagated through Monte Carlo … - Test samples [not-assessed] req3/ach1 AC=yes :: Blood analog fluid samples characterized for viscosity and density; inter-batch variability bounded and reported. - Equivalency of input parameters [not-assessed] req3/ach1 AC=yes :: Demonstration that input parameters used in validation simulations are equivalent to those in the COU simulations within stated to… | Primary hemolysis index validation against FDA pump bench data; Intermediate shear-stress distribution validation result; Monte Carlo UQ convergence validation result | (none declared) | Despite multiple credibility factor gaps, the assessment team accepted the package based on the primary hemolysis validation result and Monte Carlo UQ convergence evidence. The intermediate shear-stress distribution validation result (intermediate-orphan) was included in scope but its provenance was not fully traced. All 13 factors were reviewed; 7 assessed factors show achieved levels below requi… | judge_ready_bundle/packages/adv-2026-p2-020-w-prov-01_low_morrison-cou2-v15.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 0.7 | The property the rule checks is structurally present and machine-readable; the defeater was planted correctly; and a one-line seed change detects it. That's a rule scoping defect, not a schema gap. |  | W-PROV-01 (Critical) — Chain terminates at a node with no upstream edge and not marked isFoundationalEvidence. [Provenance] Look at: validation_results / evidence chain: any node that is neither foundational-marked nor derived from anything? Scope note from the mutation work: the rule walks upstream FROM claims, so nodes outside any claim's derivation subtree are out of scope; an "orphan" not reachable from a claim doesn't fire, and that is per-spec behavior. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
65 | adv-2026-p2-010-w-al-02_medium_nagaraja-cou1-v02 | confirm_existing | W-AL-02 | structural | khakzad/aleatoric/variability | medium | nagaraja-cou1 | COV-WRONG | W-AL-02 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: FEA software used under documented quality management system with version control and audit trail. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems with known analytical solutions solved; residuals and round-off errors documented. - Discretization error [assessed] req2/ach2 AC=yes :: Mesh convergence study demonstrating QoI change < 2% between successive refinements. - Numerical solver error [assessed] req2/ach2 AC=yes :: Iterative solver convergence criteria documented; solution verified against direct solver for representative load case. - Use error [assessed] req2/ach2 AC=yes :: Documented modeling protocol with independent peer review of model setup. - Model form [assessed] req3/ach3 AC=yes :: Continuum mechanics assumptions justified; contact formulation and material constitutive model appropriate for the loading regime. - Model inputs [assessed] req3/ach3 AC=yes :: Material properties sourced from traceable coupon testing; geometric inputs from CMM-verified CAD; input uncertainty quantified vi… - Test samples [assessed] req3/ach3 AC=yes :: N >= 10 test samples from production lots with dimensional inspection confirming conformance to design specification. - Test conditions [assessed] req3/ach2 AC=yes :: Multiple test configurations (e.g., varying offset distances or loading rates) to bracket COU loading space. - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: FEA inputs (geometry, material properties, boundary conditions) traceable to as-tested specimen measurements. - Output comparison [assessed] req3/ach3 AC=yes :: Quantitative metric (E/uval) computed for each QoI; all E/uval < 1.96 and percent difference < 10%. - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs directly correspond to regulatory submission metrics and clinical failure modes. - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario geometry, loading, and boundary conditions match or bound the COU scenario. | Construct stiffness comparison: FEA vs. bench (N=11); Yield force comparison: FEA vs. bench (N=11); Force at 40 mm displacement comparison: FEA vs. bench (N=11) | (none declared) | All 13 V&V 40 credibility factors assessed. Twelve of thirteen factors meet or exceed required levels. One factor (Test conditions, required level 3, achieved level 2) is below the required level due to single test configuration; this shortfall is accepted given that the ASTM F1717 standard mandates the executed configuration and model influence is Medium. Uncertainty quantification via Monte Carl… | judge_ready_bundle/packages/adv-2026-p2-010-w-al-02_medium_nagaraja-cou1-v02.jsonld | CORRECT-DETECTION |  | 1 | sensitivity gap — UQ declared but no hasSensitivityAnalysis link. | recorded evidence predates the 3cd9a5ff rewrite | W-AL-02 (Medium) — UQ reported but no sensitivity analysis documented. [Aleatory] Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. If the package has no UQ at all, this rule should NOT fire (W-AL-01 territory instead). | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
66 | adv-2026-p2-014-w-si-02_medium_morrison-cou1-v09 | confirm_existing | W-SI-02 | structural | khakzad/argumentation/undermining-d1 | medium | morrison-cou1 | COV-WRONG | W-SI-02 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Commercial CFD solver with documented verification test suite; version control and release notes available. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems demonstrating correct implementation of governing equations for incompressible viscous flow. - Discretization error [assessed] req2/ach2 AC=yes :: Grid convergence index (GCI) < 5% for hemolysis index QoI across all five operating points. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental boundary conditions replicated in simulation to within measurement uncertainty bounds. - Output comparison [assessed] req2/ach2 AC=yes :: Predicted RH index within 15% of experimental measurements across all five flow-speed combinations. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs directly map to hemolysis index used in regulatory submission decision criterion. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry and operating conditions sufficiently representative of the COU device and intended operating range. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this factor. | Relative hemolysis index comparison against predicate device experimental data; Mesh convergence study for CFD hemolysis prediction across 5 operating points; Pressure-flow characteristic validation against FDA pump experimental measurements | (none declared) | All seven assessed credibility factors meet or exceed required levels for MRL 2. Six factors appropriately scoped as not-assessed per ASME V&V 40-2018 MRL 2 requirements. CFD model demonstrates sufficient credibility to support identification of worst-case hemolysis operating conditions for the Class II CPB device. Relative hemolysis index predictions within 15% acceptance threshold across all fiv… | judge_ready_bundle/packages/adv-2026-p2-014-w-si-02_medium_morrison-cou1-v09.jsonld | GENERATOR-ARTIFACT |  | 1 | bindsRequirement is not missing |  | W-SI-02 (High) — Missing bindsRequirement. [Structural] Same mandatory-field logic as W-SI-01. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
67 | adv-2026-p2-004-w-ar-04_high_morrison-cou1-v05 | confirm_existing | W-AR-04 | D4 | jarzebowicz-wardzinski/argument_defeaters/D4-undercutting-model | high | morrison-cou1 | COV-WRONG | W-AR-04 | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Version-controlled solver with documented build procedures and regression test suite covering hemolysis sub-model. - Numerical code verification [assessed] req2/ach2 AC=yes :: Solver passes method of manufactured solutions (MMS) test for the Navier-Stokes equations and the scalar transport equation used i… - Discretization error [assessed] req2/ach3 AC=yes :: GCI on the primary QoI (volume-averaged shear stress scalar) must be below 5% on the production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental test conditions (fluid properties, rotational speed, flow rate) must be documented with measurement uncertainty and m… - Output comparison [assessed] req2/ach2 AC=yes :: Predicted RH index must agree with experimental RH within combined model and experimental uncertainty at all five operating points… - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: The QoI used in validation (RH index) must be the same quantity used in the COU decision criterion. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, fluid, and operating conditions must be sufficiently similar to the COU device to support extrapolation of va… - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 beyond code verification already assessed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; analyst training records on file. - Model form [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; model form uncertainty acknowledged in output comparison uncertainty budget. - Model inputs [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; input parameters documented in simulation report. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; sample provenance documented in FDA benchmark dataset metadata. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; equivalency between validation and COU inputs qualitatively confirmed. | Relative hemolysis index benchmark validation — FDA pump geometry; Mesh sensitivity and discretization error quantification | (none declared) | All seven assessed credibility factors meet or exceed their required levels for MRL 2. The six not-assessed factors are appropriately scoped out at MRL 2 per ASME V&V 40-2018 Table 3. Validation evidence demonstrates that the CFD model predicts the relative hemolysis index within combined uncertainty bounds across the full CPB operating envelope. The primary validation study was executed under bui… | judge_ready_bundle/packages/adv-2026-p2-004-w-ar-04_high_morrison-cou1-v05.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | currentModelVersion='ansys-cfx-15-build-2026-q1-build-7'; validation generated by ['ansys-cfx-15-build-2026-q1-build-7', 'ansys-cfx-15-build-2026-q1-build-7-rc2'] — 1 MISMATCH |  | W-AR-04 (High) — Validated at modelVersion ≠ currentModelVersion. [Argumentation] Look at: does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
68 | adv-2026-p2-209-nc-clean-low-confidence-but-documented_medium_nagaraja-cou1-v01 | negative_control | (not recorded) |  | control/none | medium | nagaraja-cou1 | COV-WRONG | (not recorded) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-02; W-AR-02; W-EP-01; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: https://uofa.net/synth/criteria/sqa-level2 - Numerical code verification [assessed] req2/ach2 AC=yes :: https://uofa.net/synth/criteria/code-verification-level2 - Discretization error [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/discretization-level3 - Numerical solver error [assessed] req2/ach2 AC=yes :: https://uofa.net/synth/criteria/solver-error-level2 - Use error [assessed] req2/ach2 AC=yes :: https://uofa.net/synth/criteria/use-error-level2 - Model form [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/model-form-level3 - Model inputs [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/model-inputs-level3 - Test samples [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/test-samples-level3 - Test conditions [assessed] req3/ach2 AC=yes :: https://uofa.net/synth/criteria/test-conditions-level3 - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/equivalency-inputs-level3 - Output comparison [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/output-comparison-level3 - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/qoi-relevance-level3 - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: https://uofa.net/synth/criteria/validation-relevance-level3 | Construct stiffness QoI comparison — non-cannulated; Yield force QoI comparison — non-cannulated; Force at 40 mm displacement QoI comparison — non-cannulated | (none declared) | The credibility assessment for the non-cannulated 7.5 mm pedicle screw FEA model is ACCEPTED with documented residual uncertainty. All three QoIs satisfy the quantitative acceptance criterion (E/uval < 1.96): stiffness 0.91, yield force 1.44, force-at-40mm 1.12. All FEA predictions fall within the 10% relative error acceptance band. Eleven of thirteen V&V 40 credibility factors achieve their requi… | judge_ready_bundle/packages/adv-2026-p2-209-nc-clean-low-confidence-but-documented_medium_nagaraja-cou1-v01.jsonld | GENERATOR-ARTIFACT |  | 1 | W-ON-02 fired because COU has neither applicability constraint nor operating envelope. |  | NEGATIVE CONTROL — clean package. Clean packages. Any firing beyond zero criticals is the question: is the fired-on absence genuinely present (then the package isn't clean and it's GENERATOR-ARTIFACT territory) or is the rule misreading (EXISTING-RULE-MISBEHAVIOR)? | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-02 (Medium) — UQ reported but no sensitivity analysis documented. Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
69 | adv-2026-p2-016-w-con-02_low_nagaraja-cou1-v07 | confirm_existing | W-CON-02 | structural | gohar/requirements/incorrect | low | nagaraja-cou1 | COV-HIT | W-CON-02 | TRUE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-02; W-AR-05; W-CON-02; W-EP-01; W-EP-02; W-ON-02 | Nagaraja COU1 — Non-cannulated pedicle screw compression-bending credibility assessment | V&V 40 credibility assessment for FEA prediction of ASTM F1717 static compression-bending performance of a non-cannulated 7.5 mm pedicle screw construct. COU1 is the validation activity itself — the FEA model ensemble (FDA, Ansys, DePuy Synthes) is compared against N=11 experimental constructs acros… | ProfileMinimal | Accepted | Medium | 3 | Medium | Medium | Demonstrate that the FEA modeling framework reproduces experimental compression-bending behavior of a non-cannulated pedicle screw construct within pre-established acceptance bounds (<10% QoI difference, E/uval < 1.96). | 13 | (none) | (none) | Test conditions(req3>ach2) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Ansys Mechanical release validation documentation and FDA in-house code verification records available and reviewed. - Numerical code verification [assessed] req2/ach2 AC=yes :: Benchmark problems with known analytical solutions solved; results within 1% of analytical values. - Discretization error [assessed] req3/ach3 AC=yes :: Mesh convergence study demonstrates <2% change in QoIs between medium and fine mesh; Richardson extrapolation applied. - Numerical solver error [assessed] req2/ach2 AC=yes :: Solver residuals below 1e-6 for all load steps; energy balance monitored throughout simulation. - Use error [assessed] req2/ach2 AC=yes :: Independent review of model setup by second analyst; simulation plan reviewed prior to execution. - Model form [assessed] req3/ach3 AC=yes :: Continuum FEA with nonlinear geometry and elastic-plastic material model justified by comparison to experimental failure modes. - Model inputs [assessed] req3/ach3 AC=yes :: Material properties sourced from coupon testing of same screw alloy lot; geometric dimensions from CMM measurements. - Test samples [assessed] req3/ach3 AC=yes :: N=11 constructs from production lot; sample-to-sample variability characterized. - Test conditions [assessed] req3/ach2 AC=yes :: Multiple loading rates and fixture configurations tested to characterize sensitivity of QoIs to test conditions. - Equivalency of input parameters [assessed] req3/ach3 AC=yes :: FEA boundary conditions and material inputs traceable to measured experimental values; differences documented and justified. - Output comparison [assessed] req3/ach3 AC=yes :: Percent difference <10% for all QoIs; E/uval <1.96 for all ensemble members. - Relevance of the quantities of interest [assessed] req3/ach3 AC=yes :: QoIs directly linked to clinical failure modes and regulatory submission requirements. - Relevance of the validation activities to the COU [assessed] req3/ach3 AC=yes :: Validation scenario matches COU geometry, loading, and boundary conditions; any differences documented. | Construct stiffness QoI comparison — FEA vs. experiment; Yield force QoI comparison — FEA vs. experiment; Force at 40 mm displacement QoI comparison — FEA vs. experiment | (none declared) | Twelve of thirteen V&V 40 credibility factors meet or exceed required levels for a medium model risk assessment. The single deficiency (Test conditions, achieved level 2 vs. required level 3) is acknowledged and accepted based on the conservative nature of the single test condition evaluated and the strong agreement between FEA and experimental QoIs across all other factors. Overall credibility su… | judge_ready_bundle/packages/adv-2026-p2-016-w-con-02_low_nagaraja-cou1-v07.jsonld | CORRECT-DETECTION |  | 1 | referencesIdentifier: fake-paper[DANGLES] |  | W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. [Consistency] Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-CON-02 (Medium) — Referenced identifier doesn't resolve, no external-fetch hint. Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
70 | adv-2026-p2-004-w-ar-04_high_morrison-cou2-v09 | confirm_existing | W-AR-04 | D4 | jarzebowicz-wardzinski/argument_defeaters/D4-undercutting-model | high | morrison-cou2 | COV-WRONG | W-AR-04 | FALSE | COMPOUND-01; W-AL-01; W-AR-02; W-AR-05; W-EP-01; W-EP-04; W-ON-02 | Morrison COU2 — VAD hemolysis credibility assessment | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under ventricular assist device (VAD) conditions. COU2 evaluates whether the computational model can predict absolute hemolysis levels for a Class III VAD device at Model Risk Level 5. Higher model r… | ProfileMinimal | Accepted | Low | 5 | High | High | Use CFD to predict absolute hemolysis index values for a Class III VAD device across 5 flow-speed combinations. Results must demonstrate hemolysis below absolute clinical threshold (not relative to predicate). Requires Monte Carlo uncertain… | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Model form(not-assessed); Model inputs(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | Software quality assurance(req4>ach3); Numerical code verification(req4>ach3); Discretization error(req4>ach3); Test conditions(req4>ach3); Output comparison(req4>ach3); Relevance of the quantities of interest(req5>ach3); Relevance of the validation activities to the COU(req5>ach3); Numerical solver error(req4>ach1); Use error(req4>ach1); Model form(req5>ach1); Model inputs(req5>ach1); Test samples(req4>ach1); Equivalency of input parameters(req4>ach1) | - Software quality assurance [assessed] req4/ach3 AC=yes :: Solver must be validated under a formally released build with traceable version control records and regression test suite results. - Numerical code verification [assessed] req4/ach3 AC=yes :: Order-of-accuracy studies and manufactured-solution tests executed under the production solver release. - Discretization error [assessed] req4/ach3 AC=yes :: Grid-convergence index (GCI) < 5% on NIH at finest mesh level; Richardson extrapolation performed. - Test conditions [assessed] req4/ach3 AC=yes :: In-vitro test conditions must match COU operating envelope within specified tolerances; temperature, hematocrit, and flow rate doc… - Output comparison [assessed] req4/ach3 AC=yes :: Quantitative comparison metric (e.g., E_rms) computed with experimental uncertainty bounds; all predictions within acceptance band… - Relevance of the quantities of interest [assessed] req5/ach3 AC=yes :: QoI must be the absolute NIH value directly used in the regulatory decision; surrogate quantities not acceptable at MRL 5. - Relevance of the validation activities to the COU [assessed] req5/ach3 AC=yes :: Validation geometry, fluid properties, and operating conditions must be representative of the COU device and patient population. - Numerical solver error [not-assessed] req4/ach1 AC=yes :: Iterative convergence, round-off error, and solver residual norms must be quantified and shown to be negligible relative to discre… - Use error [not-assessed] req4/ach1 AC=yes :: Independent review of simulation setup, boundary conditions, and post-processing scripts by a qualified engineer not involved in t… - Model form [not-assessed] req5/ach1 AC=yes :: Justification for selection of SST k-omega turbulence model and power-law hemolysis model for VAD conditions; sensitivity to model… - Model inputs [not-assessed] req5/ach1 AC=yes :: All uncertain model inputs identified and characterised with probability distributions; Monte Carlo inputs traceable to patient po… - Test samples [not-assessed] req4/ach1 AC=yes :: Number of experimental replicates sufficient to characterise inter-sample variability; statistical power analysis documented. - Equivalency of input parameters [not-assessed] req4/ach1 AC=yes :: Mapping between in-vitro test inputs and COU patient physiology inputs formally documented and justified. | Absolute hemolysis index validation — FDA pump 5-point sweep; Monte Carlo UQ propagation — hemolysis NIH uncertainty bands | (none declared) | The credibility assessment team accepted the simulation evidence on the basis that the validation results demonstrate NIH predictions within 12% of experimental values across all operating points and that Monte Carlo UQ confirms the 95th-percentile NIH remains below the clinical threshold. The team noted that validation evidence was generated under build 2026-q1-build-7-rc2 rather than the release… | judge_ready_bundle/packages/adv-2026-p2-004-w-ar-04_high_morrison-cou2-v09.jsonld | EXISTING-RULE-MISBEHAVIOR |  | 1 | currentModelVersion='ansys-cfx-15-build-2026-q1-build-7'; validation generated by ['ansys-cfx-15-build-2026-q1-build-7-rc2'] — 2 MISMATCH |  | W-AR-04 (High) — Validated at modelVersion ≠ currentModelVersion. [Argumentation] Look at: does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-02 (Critical) — achievedLevel < requiredLevel, decision Accepted, no offset rationale. Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-04 (High) — Factor not assessed while model_risk_level > 2. Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
71 | adv-2026-p2-113-faults-software_low-v05 | gap_probe | (none - gap_probe: no rule expected) |  | gohar/contextual/faults-software | low |  | COV-WRONG | (none - gap_probe) | FALSE | COMPOUND-01; COMPOUND-03; W-AL-01; W-AR-05; W-EP-01; W-EP-02; W-ON-02 | Morrison COU1 — CPB hemolysis credibility assessment (software-fault gap) | V&V 40 credibility assessment for CFD hemolysis prediction of the FDA generic centrifugal blood pump under cardiopulmonary bypass (CPB) conditions. COU1 evaluates whether the computational model can identify worst-case hemolysis operating conditions for a Class II CPB device at Model Risk Level 2. F… | ProfileMinimal | Accepted | Medium | 2 | High | Medium | Use CFD to identify worst-case operating conditions for hemolysis across 5 flow-speed combinations for a Class II CPB device. Results compared against predicate device hemolysis using relative hemolysis index (RH < 1 acceptance criterion). | 13 | (none) | Numerical solver error(not-assessed); Use error(not-assessed); Test samples(not-assessed); Equivalency of input parameters(not-assessed) | (none) | - Software quality assurance [assessed] req2/ach2 AC=yes :: Solver codebase under version control; documented build and test procedures; regression test suite executed prior to study. - Numerical code verification [assessed] req2/ach2 AC=yes :: Observed order of accuracy within 10% of theoretical order for primary QoI solver. - Discretization error [assessed] req2/ach2 AC=yes :: GCI < 5% for hemolysis index QoI on the production mesh. - Test conditions [assessed] req2/ach2 AC=yes :: Experimental flow rates and rotational speeds reproduced within ±3% in simulation boundary conditions. - Output comparison [assessed] req2/ach2 AC=yes :: CFD-predicted NIH within 15% of experimental NIH at all 5 operating points. - Relevance of the quantities of interest [assessed] req2/ach2 AC=yes :: QoIs used in validation directly correspond to clinical hemolysis metrics used in the COU decision. - Relevance of the validation activities to the COU [assessed] req2/ach2 AC=yes :: Validation geometry, flow regime, and operating conditions overlap with COU operating envelope. - Numerical solver error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 beyond code verification activities already performed. - Use error [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 for this COU. - Model form [assessed] req2/ach2 AC=yes :: Turbulence model and hemolysis power-law model selection justified by literature and sensitivity analysis. - Model inputs [assessed] req2/ach2 AC=yes :: Key input parameters identified and uncertainty ranges documented; sensitivity analysis performed. - Test samples [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2; FDA benchmark dataset provides sufficient experimental basis. - Equivalency of input parameters [not-assessed] req1/ach1 AC=yes :: Not required at MRL 2 given direct correspondence between validation and COU boundary conditions. | Mesh convergence study — nominal completion; Experimental hemolysis comparison — nominal completion; Method of manufactured solutions verification — nominal completion; Input parameter sensitivity analysis — nominal completion | (none declared) | All 7 assessed credibility factors meet or exceed required levels for MRL 2. The CFD hemolysis model demonstrates adequate mesh convergence (GCI 2.1%), code verification (observed order 1.97), and experimental agreement (max error 11.7%) across all 5 operating conditions. RH < 1 at all operating points confirms the device meets the hemolysis acceptance criterion. The 6 non-assessed factors are app… | judge_ready_bundle/packages/adv-2026-p2-113-faults-software_low-v05.jsonld | OUT-OF-SCOPE |  | 1 | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. |  | GAP-PROBE CASE — no rule was expected to fire. These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then: - Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar). - Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION. - Not really instantiated → GENERATOR-ARTIFACT. - Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material. | COMPOUND-01 (Critical) — Critical and High weakeners coexist. Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. COMPOUND-03 (High) — Assurance level not Low while Critical weakeners exist. Look at: assurance_level column ≠ Low AND any Critical in rules_fired. W-AL-01 (High) — ValidationResult with no uncertainty quantification. Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). W-AR-05 (High) — ValidationResult with no comparedAgainst link. Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. W-EP-01 (Critical) — Claim with no wasDerivedFrom link to evidence. Look at: whether the package's claims link to evidence at all. W-EP-02 (High) — ValidationResult with no prov:wasGeneratedBy. Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) W-ON-02 (High) — COU has neither applicability constraint nor operating envelope. Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | A - prose-only (no property exists) | Nominal-only assurance is stated in the SQA factor's acceptanceCriteria ('passing all nominal test cases') and in ValidationResult description text. No property encodes validation conditions or fault-mode coverage, so no rule can read the limitation. Numerical solver error is not-assessed but permitted at MRL 2, and W-EP-04 is gated on MRL > 2, so engine silence is per-spec. | NONE - no Tier-1 candidate names this defeater

## Cheat sheet
Stage 4 Adjudication Cheat Sheet
Use beside the worksheet. For every case, two questions IN ORDER:
Q1: Does the package actually contain the target weakener? (Check the columns named below against target_weakener.)
Q2: If yes, did the engine respond correctly? (Compare expected_rule / target_rule_fired / rules_fired.)
Verdict vocabulary
• CORRECT-DETECTION: flaw is real, engine caught it (or correctly stayed silent on a clean package).
• GENERATOR-ARTIFACT: the package does NOT instantiate the flaw it was asked to (or instantiates it malformed). Engine behavior is moot.
• EXISTING-RULE-MISBEHAVIOR: flaw handling is wrong given what the package really contains (fired when it shouldn't, silent when it should).
• REAL-GAP: package contains a genuine weakness no current rule covers (gap_probe cases mostly).
• OUT-OF-SCOPE: the weakness is real but not the catalog's job (substantive quality, prose-only content).
The four standing traps (from this month's measured findings)
1. Prose does not count. Rules read properties, not narrative. If decision_rationale or package_description SAYS a comparator/UQ/analysis exists but no structural link appears in validation_results / factors_detail, the absence is REAL and the rule firing is CORRECT-DETECTION. (~200 archived judgments got this wrong; do not join them.)
2. Precondition present is not a false positive. Six rules have the shape "X present but Y absent." A judge calling the firing wrong BECAUSE X is present has misread the rule. If X is present and Y is absent, the rule is right.
3. Placeholders satisfy presence rules. A stub ApplicabilityConstraint or empty-ish OperatingEnvelope makes W-ON-02 correctly NOT fire, even though the content is hollow. That's a rule-scope limitation, not misbehavior. Verdict on such a case: CORRECT-DETECTION (engine behaved per spec); note the hollowness in author_notes.
4. Silence where the structure never existed is GENERATOR-ARTIFACT, not rule failure. Especially W-EP-03, W-AR-04, W-CON-03 (see below). ---
Gap-probe cases (22 rows): the different question
These packages carry a weakness the catalog was NOT asked to detect. Q1 becomes: is the described weakness genuinely present and genuinely a credibility weakness? Then:
• Present, real, and no rule covers it → REAL-GAP (this is the verdict that feeds the six Tier-1 candidates; ≥3 of 6 confirming is the success bar, but judge each case on its merits, not the bar).
• Present but actually caught by an existing rule (check rules_fired) → CORRECT-DETECTION.
• Not really instantiated → GENERATOR-ARTIFACT.
• Real weakness but substantive/prose-level, not structural → OUT-OF-SCOPE, and say why in the rationale; these are boundary-section material.
Negative controls (2 rows)
Clean packages. Any firing beyond zero criticals is the question: is the fired-on absence genuinely present (then the package isn't clean and it's GENERATOR-ARTIFACT territory) or is the rule misreading (EXISTING-RULE-MISBEHAVIOR)?
Worksheet hygiene
• author_verdict, author_confidence, author_rationale on every row; rationale cites the column evidence ("validation_results shows 3 results, none with comparedAgainst"), not vibes.
• Low confidence is data; mark it honestly.
• author_section_6_7: flag only rows bearing on the six Tier-1 candidates' defeater types.
• Anything that smells like a new trap goes in author_notes verbatim; the notes column fed three findings last time.
Per-rule guidance is on the 'Rule guide' tab, and is also carried per row in the guide_target_rule / guide_rules_fired columns of adjudication_worksheet.
Source: Stage4_Adjudication_Cheat_Sheet.md
Structural-purchase scan (derived - NOT author judgment)
Columns scan_bucket / scan_basis on the adjudication_worksheet tab. Added 2026-08-17. These are a derived annotation, not an adjudication, and they are deliberately kept out of the author_* columns.
Question asked of each of the 22 gap_probe rows: does the weakness the spec asked for have any PROPERTY a rule could read, or does it exist only in narrative? Trap 1 says rules read properties, not narrative; the gap-probe branch sends real-but-prose-level weaknesses to OUT-OF-SCOPE.
Buckets: A - prose-only (no property exists) = 17 rows. B - structural handle exists = 3 rows (12, 24, 37) where the prose test does NOT apply. C - trap 4 (structure never built) = 2 rows (17, 34) where the structure the defeater needs was never built.
Sources: the 22 package .jsonld files and dev/specs/gap_probe/*.yaml only. No judge verdicts, no queue mapping, and neither adjudication_sample_key.csv nor adjudication_queue.csv was read.
Three corpus-wide findings behind the buckets: (1) across 69 ValidationResult nodes only 2-5 carry any structured field, with inconsistent keys; (2) hasContextOfUse has 8 descriptive keys and no operatingEnvelope or applicabilityConstraint in any package; (3) hasWeakener is empty in all 22.
Does the probe point at a named Tier-1 candidate? (scan_6_7_candidate)
The judge rubric (packs/core/judge_prompts/v1.1.0.md) sets REAL-GAP as: coverage_intent = gap_probe AND 'the package content visibly instantiates the SECTION 6.7 candidate's defeater' AND no/semantically-distinct rules fired. That middle condition presumes a named candidate exists.
Of the 22 gap_probe rows, 8 point at one of the six Tier-1 candidates and 14 point at none. For the unmapped rows the rubric's REAL-GAP condition has no referent, so a judge would have to supply the defeater unaided - which is the expertise the catalog exists to codify.
Mapping is derived (assistant, 2026-08-17) from each package's adversarialProvenance.specPath and the candidate list published in ADJUDICATION_INSTRUCTIONS.md. It is not an author judgment and does not prescribe what belongs in author_section_6_7.

## Rule guide
rule | severity | category | headline | look_at | full_guidance
W-EP-01 | Critical | Epistemic | Claim with no wasDerivedFrom link to evidence. | whether the package's claims link to evidence at all. | Look at: whether the package's claims link to evidence at all. SPECIAL: this rule's guard reads a class (uofa:Claim) the schema never declares, so it CANNOT fire on a schema-conformant package. If target_rule_fired = yes on this rule, the package types its claims non-conformantly → lean GENERATOR-ARTIFACT and note it. If silent on a package whose claims genuinely lack derivation, note "unfireable as shipped" in author_notes; verdict follows Q1 honestly.
W-EP-02 | High | Epistemic | ValidationResult with no prov:wasGeneratedBy. | validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.) | Look at: validation_results column. Does each result name a generating activity? Bare-IRI / stub results with no activity = flaw present, firing correct. (Morrison COU1 fires this 3x legitimately.)
W-EP-03 | High | Epistemic | Dataset vintage predates model revision date. | does the package carry BOTH a dataset vintage date AND a model revision date, as dated fields? The chain needs result → activity → used dataset. Most generated packages never build this chain (0/180 measured). If the dates aren't structurally present, the flaw was never instantiated → GENERATOR-ARTIFACT. If both dates present and vintage < revision, firing is correct. | Look at: does the package carry BOTH a dataset vintage date AND a model revision date, as dated fields? The chain needs result → activity → used dataset. Most generated packages never build this chain (0/180 measured). If the dates aren't structurally present, the flaw was never instantiated → GENERATOR-ARTIFACT. If both dates present and vintage < revision, firing is correct.
W-EP-04 | High | Epistemic | Factor not assessed while model_risk_level > 2. | factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right. | Look at: factors_not_assessed (nonzero?) AND model_risk_level (>2?). Both conditions, mechanical. If risk ≤ 2, firing is wrong; if factors_not_assessed = 0, firing is wrong; both present, firing is right.
W-AL-01 | High | Aleatory | ValidationResult with no uncertainty quantification. | validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1). | Look at: validation_results. Per-result check: any UQ attached? Fires per offending result. Prose mention of UQ in rationale does not count (trap 1).
W-AL-02 | Medium | Aleatory | UQ reported but no sensitivity analysis documented. | Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. | Precondition-present shape (trap 2): UQ must be PRESENT and SA absent. If the package has no UQ at all, this rule should NOT fire (W-AL-01 territory instead).
W-ON-01 | Critical | Ontological | No Context of Use at all. | cou_* columns populated? If the package has any COU, firing is wrong. Note: flaw cannot exist in a schema-valid package (COU is sh:minCount mandatory), so a "clean-looking" case targeting this is likely GENERATOR-ARTIFACT (package failed validation upstream). | Look at: cou_* columns populated? If the package has any COU, firing is wrong. Note: flaw cannot exist in a schema-valid package (COU is sh:minCount mandatory), so a "clean-looking" case targeting this is likely GENERATOR-ARTIFACT (package failed validation upstream).
W-ON-02 | High | Ontological | COU has neither applicability constraint nor operating envelope. | does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful. | Look at: does the COU carry either bound? Neither → fires correctly. EITHER present (even a placeholder stub) → must not fire (trap 3). This is the rule most affected by regenerated stubs; check what's actually there, not what's meaningful.
W-SI-01 | Medium | Structural | No digital signature. | Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT. | Mechanical presence check; also sh:minCount mandatory, so same note as W-ON-01: a package genuinely missing it is non-conformant → likely GENERATOR-ARTIFACT.
W-SI-02 | High | Structural | Missing bindsRequirement. | Same mandatory-field logic as W-SI-01. | Same mandatory-field logic as W-SI-01.
W-AR-01 | Critical | Argumentation | Factor has requiredLevel but no acceptance criteria. | factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present. | Look at: factors_missing_acceptance_criteria (nonzero?) and factors_detail. Precondition: requiredLevel present.
W-AR-02 | Critical | Argumentation | achievedLevel < requiredLevel, decision Accepted, no offset rationale. | factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing. | Look at: factors_below_required_level (nonzero?), decision = Accepted, and whether factors_detail / rationale carries a STRUCTURED offset rationale (a documented property, not prose). Three-part condition; all three must hold for a correct firing.
W-AR-03 | High | Argumentation | Requirement demands verification method X; activity used activityType Y. | factors_detail / validation_results for a required method versus the activity's recorded type. Both fields must exist structurally. If activityType was never populated (common in generated packages), lean GENERATOR-ARTIFACT. | Look at: factors_detail / validation_results for a required method versus the activity's recorded type. Both fields must exist structurally. If activityType was never populated (common in generated packages), lean GENERATOR-ARTIFACT.
W-AR-04 | High | Argumentation | Validated at modelVersion ≠ currentModelVersion. | does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing. | Look at: does the package carry BOTH a currentModelVersion and a per-configuration validated version? Structures rarely generated (0/180 measured). No fields → GENERATOR-ARTIFACT. Both present and unequal → correct firing.
W-AR-05 | High | Argumentation | ValidationResult with no comparedAgainst link. | validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop. | Look at: validation_results, per result: comparator entity linked? THE prose trap rule (104 archived cases): "experimental comparators exist in the narrative" does NOT clear it. Absence of the link = correct firing, full stop.
W-CON-01 | High | Consistency | Decision Accepted but a factor has neither requiredLevel nor achievedLevel. | decision = Accepted AND any factor with both levels missing (factors_detail). Note the v0.5.14 guard: factors with status 'not-assessed' are exempt; if the only levelless factors are explicitly not-assessed, firing is wrong. | Look at: decision = Accepted AND any factor with both levels missing (factors_detail). Note the v0.5.14 guard: factors with status 'not-assessed' are exempt; if the only levelless factors are explicitly not-assessed, firing is wrong.
W-CON-02 | Medium | Consistency | Referenced identifier doesn't resolve, no external-fetch hint. | declared references versus what exists in the package graph. Mechanical dangling-reference check. | Look at: declared references versus what exists in the package graph. Mechanical dangling-reference check.
W-CON-03 | High | Consistency | Evidence timestamp later than signature timestamp. | does the package structurally carry BOTH timestamps? Rarely generated (0/180). Absent → GENERATOR-ARTIFACT. Present with evidence > signature → correct firing. | Look at: does the package structurally carry BOTH timestamps? Rarely generated (0/180). Absent → GENERATOR-ARTIFACT. Present with evidence > signature → correct firing.
W-CON-04 | Medium | Consistency | ProfileComplete but no SensitivityAnalysis declared. | profile column = Complete (precondition!) and SA absence. If profile is not Complete, firing is wrong (trap 2 shape). | Look at: profile column = Complete (precondition!) and SA absence. If profile is not Complete, firing is wrong (trap 2 shape).
W-CON-05 | High | Consistency | VerificationActivity declared but no Evidence linked via wasGeneratedBy. | any declared verification activity; does evidence point back to it? Activity with nothing generated = procedural placeholder, correct firing. | Look at: any declared verification activity; does evidence point back to it? Activity with nothing generated = procedural placeholder, correct firing.
W-PROV-01 | Critical | Provenance | Chain terminates at a node with no upstream edge and not marked isFoundationalEvidence. | validation_results / evidence chain: any node that is neither foundational-marked nor derived from anything? Scope note from the mutation work: the rule walks upstream FROM claims, so nodes outside any claim's derivation subtree are out of scope; an "orphan" not reachable from a claim doesn't fire, and that is per-spec behavior. | Look at: validation_results / evidence chain: any node that is neither foundational-marked nor derived from anything? Scope note from the mutation work: the rule walks upstream FROM claims, so nodes outside any claim's derivation subtree are out of scope; an "orphan" not reachable from a claim doesn't fire, and that is per-spec behavior.
COMPOUND-01 | Critical | Compound | Critical and High weakeners coexist. | Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. | Purely derivative: check rules_fired for at least one Critical and one High Level-1 firing. Wrong only if the underlying set doesn't support it.
COMPOUND-03 | High | Compound | Assurance level not Low while Critical weakeners exist. | assurance_level column ≠ Low AND any Critical in rules_fired. | Look at: assurance_level column ≠ Low AND any Critical in rules_fired.