## Assessment Summary
Unit of Assurance — Assessment Summary
Project Name | COU Name | COU Description | Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Assessor Name | Assessment Date | Source Document | Has UQ?
Hip-Stem FEA Credibility Assessment (ISO 7206 Bending) | FEA prediction of bending stiffness, peak stress, and hotspot ranking for cementless femoral stem under ISO 7206-4/8 loading | Abaqus-based finite element model of cementless femoral stem (sizes 1–5) used to predict bending stiffness, peak von Mises stress at the neck fillet, and relative hotspot ranking under ISO 7206-4/8 style loading. Intended for design down-select and stiffness substantiation. Not approved for in vivo fatigue life or labeling claims outside ISO-style bending. | Complete | Class III | MRL 4 | High | ASME-VV40-2018 | Ethan Miller, V&V Engineer | Date of this assessment (YYYY-MM-DD) | report.md | Yes
Legend: |  | Required (Minimal) | Required (Complete) | Optional
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on C3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on K3: [Extract] ¦ Confidence: 50%
Note on L3: [Extract] ¦ Confidence: 50%

## Instructions
UofA Excel Template — Instructions
What is this? | This workbook captures a credibility assessment for one Context of Use (COU). Fill it out, then run `uofa import` to convert it to a machine-verifiable evidence package.
One workbook = one COU | If your project has multiple COUs (e.g., different device classes or risk levels for the same model), create a separate workbook for each.
Profiles | Minimal (7 fields, lightweight audit trail) is for early-stage projects or live pipeline capture. Complete (full V&V 40 assessment) is for regulatory submissions.
Color coding | Blue cells = required for Minimal profile. Green cells = required for Complete profile. White cells = optional.
Credibility Factors | Sheet 4 is pre-populated with all 13 V&V 40 factors from Table 5-1. You do not need to assess all 13 — set Assessed = "No" or "N/A" for factors not relevant to your COU.
Dropdowns | Most constrained fields use dropdown lists. If you need a custom value not in the dropdown, type it directly (the validation will warn but not prevent entry).
After filling out | Run: uofa import your-file.xlsx --sign --key keys/your.key --check
What the tool checks | Gate 1 (SHACL): Are all required fields present? Gate 2 (Jena rules): Are there quality gaps (weakeners)? Gate 3 (Integrity): Is the hash and signature valid?
More info | https://crediblesimulation.com ¦ https://github.com/cloudronin/uofa

## _Lists (hidden)
Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Entity Type | Has UQ | Pass/Fail | Factor Level (1-5) | Factor Status | Decision Outcome | Has UQ (COU-level)
Minimal | Class I | MRL 1 | Low | ASME-VV40-2018 | Requirement | Yes | Pass | 1 | assessed | Accepted | Yes
Complete | Class II | MRL 2 | Medium | NASA-STD-7009B | Model | No | Fail | 2 | not-assessed | Not accepted | No
| Class III | MRL 3 | High | FDA-2023-CMS | Dataset |  | Inconclusive | 3 | scoped-out | Conditional
| Category A | MRL 4 |  | Custom |  |  | N/A | 4 | not-applicable
| Category B | MRL 5 | [5 empty] | 5
Category C
Category D
Category E
Other

## Model & Data
Model & Data — What you modeled and compared against
Entity Type | Name | Identifier / URI | Description | Version | Source
Requirement | ISO 7206 Bending Acceptance Targets | Stable URI or local ID | Pre-declared acceptance criteria: ±10% on peak von Mises at the gauge region, ±5% on stem stiffness, and correct ordering of the top three stress hotspots across all test specimens | Version number (optional) | Where this came from (optional)
Model | Abaqus/Standard Hip-Stem FEA Model (v1.7.2) |  | Nonlinear FEA model of cementless Ti-6Al-4V femoral stem (sizes 1–5) in PMMA potting, using C3D10 elements, bilinear hardening, NLGEOM=ON, benchmarked against NAFEMS LE10 and analytical cantilever solution
Dataset | ISO 7206-4 Physical Test Data (n=8 stems) |  | MTS 858 test results for eight femoral stems (sizes 1–5, two repeats size 3): load-deflection stiffness, strain gauge peak stress, and DIC full-field hotspot maps under ISO 7206-4 loading conditions
Dataset | Ti-6Al-4V ELI Material Property Data |  | Heat-lot tensile test data (n=12): elastic modulus 110±3 GPa, Poisson 0.34, yield 880 MPa, tangent modulus 1.2 GPa; bilinear hardening model matched coupon data to within 2% in the relevant strain band
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%

## Validation Results
Validation Results — What you found when you tested the model
Result Name | Identifier / URI | Description | Compares To | Has UQ? | UQ Method | Metric Value | Pass / Fail
Peak von Mises Stress Comparison (Primary Gauge) | ValidationResult | What was done, what was found | Comparison of FEA-predicted peak von Mises stress at the primary neck fillet gauge region against DIC/strain-gauge-derived stress from eight physical test specimens | Strain gauge and DIC measurements on MTS 858, ISO 7206-4 setup; load cell uncertainty 1.0%, DIC uncertainty 2.2% | Yes | Latin Hypercube Sampling (RSS of contributions); combined output uncertainty 7.6% on peak stress | Mean absolute percent error 6.9% (primary gauge); 8.1% (secondary gauge); normalized RMSE 6.2% across all gauges; worst case within ±10% acceptance criterion | Pass
Stem Bending Stiffness Comparison | ValidationResult |  | Comparison of FEA-predicted load-deflection stiffness (slope) against physical test measurements across eight specimens under ISO 7206-4 bending | MTS 858 load-deflection data; load cell uncertainty 1.0% | Yes | Latin Hypercube Sampling; alignment contribution ~4.1% to combined output uncertainty | Mean error 3.5%; worst case 4.7%; acceptance criterion ±5% | Pass
Stress Hotspot Ranking Comparison | ValidationResult |  | Comparison of FEA-predicted top three stress hotspot ordering against DIC full-field strain maps for all eight test specimens | DIC full-field maps identifying stress maxima locations | No |  | Correct hotspot ordering in all 8 of 8 cases (100%) | Pass
Mesh Convergence Study (GCI) | ValidationResult |  | Three-level local mesh refinement around the neck fillet to assess discretization error in peak von Mises stress and stem compliance | Richardson extrapolation / Grid Convergence Index methodology | Yes | Grid Convergence Index (GCI) at 95% confidence | GCI 1.8% on peak von Mises at notch root; stem compliance change <0.8% between last two meshes | Pass
Independent Model Reproduction Check | ValidationResult |  | Independent reviewer (S. Liang, not on design team) reproduced the size-3 model from the Git repository and compared stress and stiffness results to the primary analysis | Primary FEA results (E. Miller) | No |  | Stress within 1.1%; stiffness within 0.6% | Pass
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on E4: [Extract] ¦ Confidence: 50%
Note on F4: [Extract] ¦ Confidence: 50%
Note on G4: [Extract] ¦ Confidence: 50%
Note on H4: [Extract] ¦ Confidence: 50%
Note on I4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on H5: [Extract] ¦ Confidence: 50%
Note on I5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on G6: [Extract] ¦ Confidence: 50%
Note on H6: [Extract] ¦ Confidence: 50%
Note on I6: [Extract] ¦ Confidence: 50%
Note on A7: [Extract] ¦ Confidence: 50%
Note on B7: [Extract] ¦ Confidence: 50%
Note on D7: [Extract] ¦ Confidence: 50%
Note on E7: [Extract] ¦ Confidence: 50%
Note on F7: [Extract] ¦ Confidence: 50%
Note on H7: [Extract] ¦ Confidence: 50%
Note on I7: [Extract] ¦ Confidence: 50%

## Credibility Factors
Credibility Factors — V&V 40 Table 5-1 Assessment
Complete profile only. Assess each relevant factor: set Required Level (goal given model risk) and Achieved Level (what evidence supports).
Factor Type | Category | Required Level | Achieved Level | Acceptance Criteria | Rationale | Factor Status | Linked Evidence
V&V 40 factor (pre-filled, do not edit) | Grouping (pre-filled) | Credibility goal for this factor (1-5) | Level achieved by evidence (1-5) | Goal and basis for the required level (from V&V 40 Tables 3-4) | Brief justification (esp. if achieved < required) | assessed / not-assessed / scoped-out / not-applicable | URI of validation result that supports this factor (optional)
Software quality assurance | Verification — Code | 4 | 4 | Commercial solver with documented version control, regression testing or benchmark suite, and reproducible results | Abaqus/Standard 2023 FD03 used; model and scripts tracked in Git (repo MS-hipstem-fea, tag v1.7.2). Results hash-matched within 0.2% across Windows 11 and RHEL 8 platforms. Benchmarked against NAFEMS LE10 (0.4% displacement error) and analytical cantilever (0.3% stress error). All assumptions and metadata cataloged in Confluence FEA-HIP-ISO-BEND v1.7. | assessed
Numerical code verification | Verification — Code | 4 | 4 | Verification against analytical solutions and recognized benchmark problems demonstrating code correctness | Abaqus verified against NAFEMS LE10 bending benchmark (0.4% error in displacement) and a custom cantilever with analytical tip-load solution (0.3% error in stress). Both benchmarks are documented in the report and directly relevant to bending stress problems. | assessed
Discretization error | Verification — Calculation | 4 | 4 | Mesh convergence study with quantified GCI or equivalent; discretization contribution to uncertainty bounded | Three-level mesh refinement study performed around the neck fillet. GCI of 1.8% (95% confidence) on peak von Mises; stem compliance changed <0.8% between last two meshes. Local element size 0.3 mm at neck blend; growth rate ≤1.3. Mesh contribution isolated as ~1.8% in LHS uncertainty decomposition. | assessed
Numerical solver error | Verification — Calculation | 4 | 4 | Documented convergence criteria, residual targets met, energy balance error bounded | Nonlinear solution with NLGEOM=ON; force residual tolerance 0.5%, displacement norm 0.1 mm, achieved within 18–32 increments. Energy balance error <1.5% at convergence. Contact overclosure at potting interface averaged 2.4 μm (<0.5% of smallest element edge), confirming well-behaved constraint system. | assessed
Use error | Verification — Calculation | 4 | 4 | Independent review of model setup, boundary conditions verified against test conditions, mesh quality checks documented | Independent reviewer S. Liang (not on design team) reproduced the size-3 model from the repository, obtaining stress within 1.1% and stiffness within 0.6%. All assumptions, CAD trims, and test metadata cataloged in Confluence FEA-HIP-ISO-BEND v1.7. Boundary conditions (potting angle, embedding length) measured per specimen and used as model inputs. | assessed
Model form | Validation — Model | 4 | 3 | Governing equations and constitutive model justified; limitations documented; physics assumptions appropriate for COU | Nonlinear static FEA with bilinear hardening Ti-6Al-4V constitutive model justified against coupon data (within 2% in relevant strain band). NLGEOM=ON for fixture compliance. Bonded stem–PMMA contact justified by irreversible cure; sensitivity to friction (μ=0.3) shown to reduce peak stress by only 1.2%. Scope and limitations explicitly stated (not for fatigue, fretting, corrosion, gait cycles, or temperatures outside 23±2°C). No creep model. Model form adequate for linear-elastic dominated bending but elastoplastic regime not deeply validated beyond coupon-level. | assessed
Model inputs | Validation — Model | 4 | 4 | Material properties from physical testing; geometry from production parts; boundary conditions from measurements; input uncertainties characterized | Elastic modulus 110±3 GPa and Poisson 0.34 from heat-lot tensile tests (n=12). Bilinear hardening parameters (yield 880 MPa, tangent 1.2 GPa) matched coupon data within 2%. PMMA modulus 3.0 GPa from vendor certs. Blend radius measured on production parts (2.1±0.05 mm). Per-specimen potting angle and embedding length measured and applied as individual model inputs. Input uncertainties for E, angle, embedding length, and load magnitude quantified and propagated via LHS. | assessed
Test samples | Validation — Comparator | 4 | 3 | Adequate number of specimens covering size range; production-representative samples; statistical characterization | Eight femoral stems tested covering sizes 1–5 (two repeats for size 3). Samples are production parts. However, no explicit statistical characterization of sample population or formal sample-size justification is documented. Sizes 1, 2, 4, 5 have single repeats, limiting statistical power for those sizes. Coverage is adequate for design down-select but marginal for formal statistical claims. | assessed
Test conditions | Validation — Comparator | 4 | 4 | Calibrated instruments, controlled test environment, measurement uncertainty quantified, test conducted per recognized standard | Tests conducted per ISO 7206-4 on MTS 858. Load cell uncertainty 1.0%; DIC measurement uncertainty 2.2% (calibrated with 25 mm grid plate). Load vector 10° anterior, 10° lateral; 80 mm embedding length. Potting length and load angle measured per specimen. Temperature 23±2°C. Standard-compliant setup with documented instrument calibration. | assessed
Equivalency of input parameters | Validation — Assessment | 4 | 4 | Per-specimen model inputs matched to measured test conditions; input uncertainties propagated to outputs | Per-specimen potting angle, embedding length, and load magnitude measured and applied directly as FEA inputs. Material properties from the same production heat lots. Input uncertainty propagation via LHS sampling (E ±3 GPa, potting angle ±0.5°, embedding length ±1 mm, load ±1%) yields combined output uncertainty 7.6% RSS. No parameter fitting to bench data performed. | assessed
Output comparison | Validation — Assessment | 4 | 4 | Quantitative error metrics against pre-declared acceptance criteria; multiple comparison points; uncertainty bands reported | Pre-declared acceptance targets met: MAPE 6.9% at primary gauge (criterion ±10%); MAPE 8.1% at secondary gauge; normalized RMSE 6.2% across all gauges; stiffness error 3.5% mean / 4.7% worst case (criterion ±5%); hotspot ranking correct in 8/8 cases. Combined model output uncertainty 7.6% RSS. All eight specimens reported individually with no outliers exceeding the 10% stress tolerance. | assessed
Relevance of the quantities of interest | Applicability | 4 | 4 | QoI directly measures the safety/performance concern and is measurable both computationally and experimentally | QoIs (bending stiffness, peak von Mises stress at neck fillet, hotspot ranking) directly correspond to structural performance and failure risk under ISO 7206-4/8 bending, which is the regulatory and design substantiation concern. All three QoIs are computed by FEA and measured experimentally (load-deflection, strain gauges, DIC), providing direct comparison capability. | assessed
Relevance of the validation activities to the COU | Applicability | 4 | 3 | Validation conditions match COU operating envelope; same geometry, loading regime, and physics; gaps identified and bounded | Validation directly replicates ISO 7206-4 loading (80 mm embedding, 10°/10° load vector, 230 N·m target moment) for all five stem sizes, matching the declared COU. Scope explicitly limited to ISO 7206-4/8 bending between 150–260 N·m; gaps acknowledged (no in vivo gait, fretting, fatigue, or temperature variation beyond 23±2°C). Validation does not cover upper and lower bounding moments of the declared 150–260 N·m range explicitly — only 230 N·m nominal tested. Minor gap in full envelope coverage reduces level from 4 to 3. | assessed
Note on C5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on G5: [Extract] ¦ Confidence: 50%
Note on C6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on G6: [Extract] ¦ Confidence: 50%
Note on C7: [Extract] ¦ Confidence: 50%
Note on D7: [Extract] ¦ Confidence: 50%
Note on E7: [Extract] ¦ Confidence: 50%
Note on F7: [Extract] ¦ Confidence: 50%
Note on G7: [Extract] ¦ Confidence: 50%
Note on C8: [Extract] ¦ Confidence: 50%
Note on D8: [Extract] ¦ Confidence: 50%
Note on E8: [Extract] ¦ Confidence: 50%
Note on F8: [Extract] ¦ Confidence: 50%
Note on G8: [Extract] ¦ Confidence: 50%
Note on C9: [Extract] ¦ Confidence: 50%
Note on D9: [Extract] ¦ Confidence: 50%
Note on E9: [Extract] ¦ Confidence: 50%
Note on F9: [Extract] ¦ Confidence: 50%
Note on G9: [Extract] ¦ Confidence: 50%
Note on C10: [Extract] ¦ Confidence: 50%
Note on D10: [Extract] ¦ Confidence: 50%
Note on E10: [Extract] ¦ Confidence: 50%
Note on F10: [Extract] ¦ Confidence: 50%
Note on G10: [Extract] ¦ Confidence: 50%
Note on C11: [Extract] ¦ Confidence: 50%
Note on D11: [Extract] ¦ Confidence: 50%
Note on E11: [Extract] ¦ Confidence: 50%
Note on F11: [Extract] ¦ Confidence: 50%
Note on G11: [Extract] ¦ Confidence: 50%
Note on C12: [Extract] ¦ Confidence: 50%
Note on D12: [Extract] ¦ Confidence: 50%
Note on E12: [Extract] ¦ Confidence: 50%
Note on F12: [Extract] ¦ Confidence: 50%
Note on G12: [Extract] ¦ Confidence: 50%
Note on C13: [Extract] ¦ Confidence: 50%
Note on D13: [Extract] ¦ Confidence: 50%
Note on E13: [Extract] ¦ Confidence: 50%
Note on F13: [Extract] ¦ Confidence: 50%
Note on G13: [Extract] ¦ Confidence: 50%
Note on C14: [Extract] ¦ Confidence: 50%
Note on D14: [Extract] ¦ Confidence: 50%
Note on E14: [Extract] ¦ Confidence: 50%
Note on F14: [Extract] ¦ Confidence: 50%
Note on G14: [Extract] ¦ Confidence: 50%
Note on C15: [Extract] ¦ Confidence: 50%
Note on D15: [Extract] ¦ Confidence: 50%
Note on E15: [Extract] ¦ Confidence: 50%
Note on F15: [Extract] ¦ Confidence: 50%
Note on G15: [Extract] ¦ Confidence: 50%
Note on C16: [Extract] ¦ Confidence: 50%
Note on D16: [Extract] ¦ Confidence: 50%
Note on E16: [Extract] ¦ Confidence: 50%
Note on F16: [Extract] ¦ Confidence: 50%
Note on G16: [Extract] ¦ Confidence: 50%
Note on C17: [Extract] ¦ Confidence: 50%
Note on D17: [Extract] ¦ Confidence: 50%
Note on E17: [Extract] ¦ Confidence: 50%
Note on F17: [Extract] ¦ Confidence: 50%
Note on G17: [Extract] ¦ Confidence: 50%

## Decision
Decision — The credibility judgment
Decision Outcome | Decision Rationale | Criteria Set | Decided By | Decision Date
Accepted | The hip-stem FEA model met all pre-declared acceptance criteria: peak von Mises stress error 6.9% (criterion ±10%), stem stiffness error 3.5% mean / 4.7% worst case (criterion ±5%), and correct hotspot ordering in all 8/8 specimens. Combined output uncertainty of 7.6% RSS is within the acceptance margin. Mesh GCI of 1.8%, solver convergence confirmed, and independent reproduction by S. Liang within 1.1% stress / 0.6% stiffness provide high confidence in model fidelity. Model is accepted for design down-select, stiffness substantiation, and hotspot ranking under ISO 7206-4/8 bending for sizes 1–5 within the stated operating bounds (150–260 N·m, 23±2°C). Model is not approved for in vivo fatigue life prediction, fretting, corrosion-fatigue, or labeling claims outside ISO-style bending without further validation evidence. | Auto-filled from Assessment Summary | Ethan Miller (V&V Lead), concurred by S. Liang (MDR Reviewer)
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%