## Assessment Summary
Unit of Assurance — Assessment Summary
Project Name | COU Name | COU Description | Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Assessor Name | Assessment Date | Source Document | Has UQ?
Hip Stem Neck Stress FEA — ISO 7206-4 | Predict peak von Mises stress at the Ti-6Al-4V stem neck fillet under ISO 7206-4 bending to support design down-select and yield margin demonstration | Abaqus/Standard model of a hip stem subjected to the ISO 7206-4 cantilever bending fixture. The model predicts peak von Mises stress and surface strain at the neck fillet over the load range 2.3–3.0 kN. Physical bench tests remain the regulatory compliance evidence; the model informs geometry choices (fillet radius 2.3–2.7 mm) for Rev C–E designs ahead of test. Target: ≤10% discrepancy to strain gauge data at the critical location; numerical error <5%. | Complete | Class III | MRL 3 | Medium | ASME-VV40-2018 | A. Singh (FEA V&V Lead) | Date of this assessment (YYYY-MM-DD) | report.md | Yes
Legend: |  | Required (Minimal) | Required (Complete) | Optional
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on C3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on K3: [Extract] ¦ Confidence: 50%
Note on L3: [Extract] ¦ Confidence: 50%

## Instructions
UofA Excel Template — Instructions
What is this? | This workbook captures a credibility assessment for one Context of Use (COU). Fill it out, then run `uofa import` to convert it to a machine-verifiable evidence package.
One workbook = one COU | If your project has multiple COUs (e.g., different device classes or risk levels for the same model), create a separate workbook for each.
Profiles | Minimal (7 fields, lightweight audit trail) is for early-stage projects or live pipeline capture. Complete (full V&V 40 assessment) is for regulatory submissions.
Color coding | Blue cells = required for Minimal profile. Green cells = required for Complete profile. White cells = optional.
Credibility Factors | Sheet 4 is pre-populated with all 13 V&V 40 factors from Table 5-1. You do not need to assess all 13 — set Assessed = "No" or "N/A" for factors not relevant to your COU.
Dropdowns | Most constrained fields use dropdown lists. If you need a custom value not in the dropdown, type it directly (the validation will warn but not prevent entry).
After filling out | Run: uofa import your-file.xlsx --sign --key keys/your.key --check
What the tool checks | Gate 1 (SHACL): Are all required fields present? Gate 2 (Jena rules): Are there quality gaps (weakeners)? Gate 3 (Integrity): Is the hash and signature valid?
More info | https://crediblesimulation.com ¦ https://github.com/cloudronin/uofa

## _Lists (hidden)
Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Entity Type | Has UQ | Pass/Fail | Factor Level (1-5) | Factor Status | Decision Outcome | Has UQ (COU-level)
Minimal | Class I | MRL 1 | Low | ASME-VV40-2018 | Requirement | Yes | Pass | 1 | assessed | Accepted | Yes
Complete | Class II | MRL 2 | Medium | NASA-STD-7009B | Model | No | Fail | 2 | not-assessed | Not accepted | No
| Class III | MRL 3 | High | FDA-2023-CMS | Dataset |  | Inconclusive | 3 | scoped-out | Conditional
| Category A | MRL 4 |  | Custom |  |  | N/A | 4 | not-applicable
| Category B | MRL 5 | [5 empty] | 5
Category C
Category D
Category E
Other

## Model & Data
Model & Data — What you modeled and compared against
Entity Type | Name | Identifier / URI | Description | Version | Source
Requirement | Peak neck fillet stress and yield margin requirement | Stable URI or local ID | Model must predict peak von Mises stress at the neck fillet within ≤10% of strain gauge data and numerical error <5%; demonstrate margin to 0.2% yield (825 MPa) under ISO 7206-4 loading for fillet radii 2.3–2.7 mm | Version number (optional) | Where this came from (optional)
Model | Abaqus/Standard Hip Stem Neck FEA (v1.4.2) |  | C3D10 elastoplastic FEA model of Rev C Ti-6Al-4V hip stem in PMMA potting under ISO 7206-4 bending fixture; predicts peak von Mises stress and surface strain at the neck fillet
Dataset | ISO 7206-4 Bench Correlation Dataset (TR-2026-041) |  | Strain gauge (3 mm stacked gauges) and DIC measurements on three Rev C stems tested per ISO 7206-4 at 2.5 kN by Test Lab B; mean strain 1350 με (SD 40)
Dataset | Material Properties Dataset (ASTM E8 coupons, LAB-MAT-2026-08) |  | Ti-6Al-4V lot 6K tensile coupons (E=110.7 GPa, σ0.2=828 MPa, Voce hardening) and PMMA three-point bend coupons (E=2.8±0.2 GPa) providing model input material curves
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%

## Validation Results
Validation Results — What you found when you tested the model
Result Name | Identifier / URI | Description | Compares To | Has UQ? | UQ Method | Metric Value | Pass / Fail
Mesh Refinement / Discretization Convergence Study | ValidationResult | What was done, what was found | Three mesh levels at the neck fillet (min edge 0.2, 0.3, 0.45 mm; 0.9M, 1.6M, 2.2M elements). Richardson extrapolation of peak stress showed 1.9% change between the two finest meshes; energy-norm change 1.5%. | Richardson extrapolation / grid convergence | Yes | Richardson extrapolation; ±2% discretization uncertainty carried for QoI | 1.9% change finest two meshes; energy-norm 1.5%; carried uncertainty ±2% | Pass
Strain Gauge Correlation at 2.5 kN | ValidationResult |  | Model-predicted surface strain at gauge centroid (1405 με) compared to mean measured strain from three instrumented stems (1350 με, SD 40) tested per ISO 7206-4. Average discrepancy 4.1%; all three stems within 7%. | Bench strain gauge measurements TR-2026-041 | Yes | Sample mean and SD across 3 specimens; comparison to ≤10% acceptance criterion | Mean error 4.1%; max 7%; SD of measurements 40 με | Pass
Full-Field DIC Surface Strain Comparison | ValidationResult |  | Full-field DIC (17 px subset) surface strain over a 12×12 mm window near the fillet compared to model surface strain field. RMS error 5.3% over the comparison window. | DIC full-field strain map at 2.5 kN | No |  | RMS error 5.3% | Pass
Uncertainty Propagation and Yield Margin Assessment | ValidationResult |  | Latin hypercube sampling (n=300) over E, potting modulus, fillet radius, and indenter offset using measured distributions. 97.5th percentile peak von Mises = 612 MPa; combined with ±2% numerical uncertainty gives 624 MPa upper bound. Margin to 0.2% yield (825 MPa) = 1.32. | 0.2% yield stress 825 MPa | Yes | Latin hypercube sampling n=300; Morris screening + Sobol first-order indices; ±2% numerical component | 97.5th percentile 612 MPa; upper bound 624 MPa; margin 1.32 | Pass
Contact Penalty Sensitivity | ValidationResult |  | Indenter contact penalty stiffness varied from 5e6 to 2e7 N/mm. Local strain at QoI location shifted by <1%, confirming non-influential contact formulation choice. | Baseline 1e7 N/mm penalty result | No |  | <1% shift in local strain | Pass
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on E4: [Extract] ¦ Confidence: 50%
Note on F4: [Extract] ¦ Confidence: 50%
Note on G4: [Extract] ¦ Confidence: 50%
Note on H4: [Extract] ¦ Confidence: 50%
Note on I4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on H5: [Extract] ¦ Confidence: 50%
Note on I5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on G6: [Extract] ¦ Confidence: 50%
Note on H6: [Extract] ¦ Confidence: 50%
Note on I6: [Extract] ¦ Confidence: 50%
Note on A7: [Extract] ¦ Confidence: 50%
Note on B7: [Extract] ¦ Confidence: 50%
Note on D7: [Extract] ¦ Confidence: 50%
Note on E7: [Extract] ¦ Confidence: 50%
Note on F7: [Extract] ¦ Confidence: 50%
Note on H7: [Extract] ¦ Confidence: 50%
Note on I7: [Extract] ¦ Confidence: 50%

## Credibility Factors
Credibility Factors — V&V 40 Table 5-1 Assessment
Complete profile only. Assess each relevant factor: set Required Level (goal given model risk) and Achieved Level (what evidence supports).
Factor Type | Category | Required Level | Achieved Level | Acceptance Criteria | Rationale | Factor Status | Linked Evidence
V&V 40 factor (pre-filled, do not edit) | Grouping (pre-filled) | Credibility goal for this factor (1-5) | Level achieved by evidence (1-5) | Goal and basis for the required level (from V&V 40 Tables 3-4) | Brief justification (esp. if achieved < required) | assessed / not-assessed / scoped-out / not-applicable | URI of validation result that supports this factor (optional)
Software quality assurance | Verification — Code | 3 | 4 | Commercial solver with documented QA, version control, regression tests, independent peer review | Abaqus/Standard 2022.HF7 used on RHEL 8.6 (double precision); spot-check on Windows 11 confirmed <0.5% difference. Model files, scripts, and material tables under Git version control (repo mdx-hipstem, tag v1.4.2); all runs recorded with input file hash and solver version. Automated cantilever patch test (<0.02% error) and manufactured bending field on a tapered bar (<0.5%) performed. Peer review by R. Chen (2026-07-18) with five findings addressed; sign-off captured in QA-REC-2026-19. | assessed
Numerical code verification | Verification — Code | 3 | 4 | Comparison to analytical/manufactured solutions demonstrating code solves governing equations correctly | Automated sanity test against linear cantilever patch field (<0.02% error) and method of manufactured solutions on a tapered bar in bending (<0.5% error) are documented. These confirm correct implementation of elastoplastic solver for the element type used. Results recorded in Git repository. | assessed
Discretization error | Verification — Calculation | 3 | 4 | Mesh convergence study with quantified numerical uncertainty <5% for QoI | Three mesh levels focused at the 2.5 mm neck radius (min edge 0.2/0.3/0.45 mm; 0.9M/1.6M/2.2M C3D10 elements). Peak stress changed 1.9% between the two finest levels; energy-norm change 1.5%. Discretization uncertainty formally carried as ±2% for the QoI, satisfying the <5% numerical error target. Time-step sensitivity (step subdivision) showed <0.3% change. | assessed
Numerical solver error | Verification — Calculation | 3 | 4 | Iterative solver convergence demonstrated with residual targets documented | Nonlinear load ramped in 20 increments; force residual <0.1% of reference load at convergence for each step. Contact stabilization disabled except for the initial settle step. Step subdivision sensitivity confirmed <0.3% change in QoI, demonstrating solver convergence is not limiting accuracy. | assessed
Use error | Verification — Calculation | 3 | 4 | Independent review of model setup, BCs, mesh quality, and post-processing; documented findings addressed | Peer review by R. Chen (independent, not an author) on 2026-07-18 covered contact penalty range, time incrementation, metrology use, mesh transition, and DIC alignment — five findings all addressed. Execution record and sign-off in QA-REC-2026-19. Two senior FEA engineers (8 and 12 yrs experience) executed the model; lab liaison confirmed fixture details. Boundary conditions (potting depth, load point, indenter geometry) verified against ISO 7206-4 and physical measurements. | assessed
Model form | Validation — Model | 3 | 3 | Governing equations justified for the physics; key assumptions documented with supporting rationale | Continuum elastoplastic model with isotropic Voce hardening selected for monotonic bending of Ti-6Al-4V; large-deformation on; C3D10 elements appropriate for stress concentrations. Key assumptions documented: no microstructural anisotropy, machining residual stresses neglected (metrology showed no tensile bias), temperature effects ignored (lab 22±1°C), contact formulation shown non-influential. Limitations explicitly scoped out (fatigue, fretting, gait spectra). Justification is qualitative; quantified model-form uncertainty not separately isolated beyond the validation comparison. | assessed
Model inputs | Validation — Model | 3 | 4 | Material properties from traceable coupon testing; geometry from metrology; boundary conditions from measurements; input uncertainties characterized | Ti-6Al-4V E=110.7 GPa and σ0.2=828 MPa from ASTM E8 coupons (lot 6K); Voce hardening curves digitized and fit. PMMA potting modulus E=2.8±0.2 GPa from three-point bend coupons (LAB-MAT-2026-08), not reverse-fitted from stem runs. Geometry from NX-2026-06 signed CAD; fillet radius confirmed by metrology on tested parts (2.48–2.52 mm). Indenter radius and fixture geometry verified by calipers; machine alignment within 0.3°. All source reports linked in EDRMS with revision locks. Input uncertainty distributions used in LHS propagation (E, potting modulus, fillet radius, indenter offset). | assessed
Test samples | Validation — Comparator | 3 | 3 | Adequate number of production-representative test articles with statistical characterization | Three Rev C stems instrumented with stacked 3 mm strain gauges tested per ISO 7206-4 (TR-2026-041). Sample size is modest (n=3) but consistent with ISO 7206-4 practice for design-phase correlation. SD of measured strain (40 με out of 1350 με mean, ~3%) indicates good repeatability. Samples are from the same revision as the model geometry. No explicit sample-size justification or power analysis documented. | assessed
Test conditions | Validation — Comparator | 3 | 3 | Calibrated instruments, controlled environment, measurement uncertainty quantified, test standard compliance | Tests performed per ISO 7206-4 by Test Lab B (TR-2026-041). Indenter hardness/radius confirmed; machine alignment within 0.3°; lab temperature 22±1°C. DIC subset 17 px used for full-field comparison. Strain gauge type and calibration mentioned (stacked 3 mm gauges) but formal measurement uncertainty budget for gauges not explicitly stated. Controlled conditions adequately documented for design-phase use. | assessed
Equivalency of input parameters | Validation — Assessment | 3 | 4 | Model inputs match experimental conditions; measurement uncertainty propagated; boundary condition matching demonstrated | Validation runs used metrology-confirmed fillet radius (2.48–2.52 mm) rather than nominal CAD value. Potting modulus set from separate coupon tests (not tuned to stem runs). Load (2.5 kN), potting depth (40 mm per ISO), indenter geometry, and fixture alignment all matched to test conditions. Indenter contact penalty sensitivity confirmed non-influential (<1% shift). Sobol sensitivity analysis identifies key input drivers; LHS propagation quantifies combined effect of input uncertainties on QoI. | assessed
Output comparison | Validation — Assessment | 3 | 4 | Quantitative error metrics against experimental data at multiple points; within acceptance criterion; uncertainty bands reported | Strain gauge comparison at gauge centroid: mean error 4.1%, max 7% across three stems — within the ≤10% acceptance criterion. Full-field DIC RMS error 5.3% over 12×12 mm window provides spatially distributed validation. Uncertainty propagation yields 97.5th-percentile stress with combined numerical uncertainty (624 MPa), giving margin 1.32 to yield. No tuning of neck behavior to test data. Multiple comparison points (point gauge + full-field DIC) and quantified uncertainty bands provided. | assessed
Relevance of the quantities of interest | Applicability | 3 | 4 | QoI directly addresses the safety/performance concern and is measurable both computationally and experimentally | Peak von Mises stress at the neck fillet is the direct mechanical quantity governing yield margin and design down-select — the stated purpose of the model. Surface strain at the gauge centroid is the measurable experimental analog used for validation. DIC provides additional spatial coverage. The QoI is computed directly from the FEA and measured directly in the bench test, with no indirect inference required. | assessed
Relevance of the validation activities to the COU | Applicability | 3 | 4 | Validation conditions match COU geometry, loading, physics regime, and operating range; gaps documented | Validation performed on Rev C stems under ISO 7206-4 fixture at 2.5 kN — the same fixture, load, geometry, and material as the COU. Fillet radius in validation (2.48–2.52 mm) is within the parameterized design range (2.3–2.7 mm). Limitations and out-of-scope uses explicitly documented (no fatigue, no patient gait, no fretting, no temperature excursions, no extrapolation beyond 2.3–2.7 mm fillet). Load range 2.3–3.0 kN assessed. Gap: validation at a single load (2.5 kN) rather than across the full 2.3–3.0 kN range. | assessed
Note on C5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on G5: [Extract] ¦ Confidence: 50%
Note on C6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on G6: [Extract] ¦ Confidence: 50%
Note on C7: [Extract] ¦ Confidence: 50%
Note on D7: [Extract] ¦ Confidence: 50%
Note on E7: [Extract] ¦ Confidence: 50%
Note on F7: [Extract] ¦ Confidence: 50%
Note on G7: [Extract] ¦ Confidence: 50%
Note on C8: [Extract] ¦ Confidence: 50%
Note on D8: [Extract] ¦ Confidence: 50%
Note on E8: [Extract] ¦ Confidence: 50%
Note on F8: [Extract] ¦ Confidence: 50%
Note on G8: [Extract] ¦ Confidence: 50%
Note on C9: [Extract] ¦ Confidence: 50%
Note on D9: [Extract] ¦ Confidence: 50%
Note on E9: [Extract] ¦ Confidence: 50%
Note on F9: [Extract] ¦ Confidence: 50%
Note on G9: [Extract] ¦ Confidence: 50%
Note on C10: [Extract] ¦ Confidence: 50%
Note on D10: [Extract] ¦ Confidence: 50%
Note on E10: [Extract] ¦ Confidence: 50%
Note on F10: [Extract] ¦ Confidence: 50%
Note on G10: [Extract] ¦ Confidence: 50%
Note on C11: [Extract] ¦ Confidence: 50%
Note on D11: [Extract] ¦ Confidence: 50%
Note on E11: [Extract] ¦ Confidence: 50%
Note on F11: [Extract] ¦ Confidence: 50%
Note on G11: [Extract] ¦ Confidence: 50%
Note on C12: [Extract] ¦ Confidence: 50%
Note on D12: [Extract] ¦ Confidence: 50%
Note on E12: [Extract] ¦ Confidence: 50%
Note on F12: [Extract] ¦ Confidence: 50%
Note on G12: [Extract] ¦ Confidence: 50%
Note on C13: [Extract] ¦ Confidence: 50%
Note on D13: [Extract] ¦ Confidence: 50%
Note on E13: [Extract] ¦ Confidence: 50%
Note on F13: [Extract] ¦ Confidence: 50%
Note on G13: [Extract] ¦ Confidence: 50%
Note on C14: [Extract] ¦ Confidence: 50%
Note on D14: [Extract] ¦ Confidence: 50%
Note on E14: [Extract] ¦ Confidence: 50%
Note on F14: [Extract] ¦ Confidence: 50%
Note on G14: [Extract] ¦ Confidence: 50%
Note on C15: [Extract] ¦ Confidence: 50%
Note on D15: [Extract] ¦ Confidence: 50%
Note on E15: [Extract] ¦ Confidence: 50%
Note on F15: [Extract] ¦ Confidence: 50%
Note on G15: [Extract] ¦ Confidence: 50%
Note on C16: [Extract] ¦ Confidence: 50%
Note on D16: [Extract] ¦ Confidence: 50%
Note on E16: [Extract] ¦ Confidence: 50%
Note on F16: [Extract] ¦ Confidence: 50%
Note on G16: [Extract] ¦ Confidence: 50%
Note on C17: [Extract] ¦ Confidence: 50%
Note on D17: [Extract] ¦ Confidence: 50%
Note on E17: [Extract] ¦ Confidence: 50%
Note on F17: [Extract] ¦ Confidence: 50%
Note on G17: [Extract] ¦ Confidence: 50%

## Decision
Decision — The credibility judgment
Decision Outcome | Decision Rationale | Criteria Set | Decided By | Decision Date
Accepted | All 13 credibility factors meet or exceed their required levels. The model achieves ≤10% discrepancy to strain gauge data (mean 4.1%, max 7%) and <5% numerical error (±2% discretization uncertainty carried). Uncertainty propagation demonstrates a margin of 1.32 to 0.2% yield at the 97.5th percentile. Independent peer review completed (QA-REC-2026-19). Acceptance is conditioned on: (1) using a mesh no coarser than 0.3 mm local size at the neck fillet; (2) Abaqus/Standard 2022.HF7 or later with identical material curves; (3) application restricted to predicting neck fillet stresses under ISO 7206-4 fixture and load range 2.3–3.0 kN for Rev C–E stems with measured fillet radii 2.3–2.7 mm; (4) model not to be used for fatigue life estimation, patient gait spectra, fretting/wear, or non-ISO fixtures. | Auto-filled from Assessment Summary | A. Singh (FEA V&V Lead); concurred by J. Patel (Chief Engineer) | 2026-08-06
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%